## Sheet1
Test# | Application | Block | Paramter | Test Type | Test Type 2 | Procedure | Test Setup | Methodology | Priority | Pass Criteria | Status | Ownership | Comments | Test Level
1 | System | Uart<->USB | Communication | Functional | Self Test | Send and receive data via UART | Diagram 1 | Send and receive a specific word by Uart | L | returns same word |  |  |  | Final Functional
2 | System | nRf52840 | HF XTAL, PLL and LF OSC | Functional | Self Test | Set the internal Clock and toggle a GPIO in order to collect data on oscilloscope | Diagram 1 | Change to other available frequency and repeat the test | L | Clocks are stable |  |  | Get realibility report from supplier. Can crosscheck between crystal oscillators | Charaterization
4 | System | Memory | Enable/Disable | Functional | Self Test | Check the access to memory when enabled and disabled | Diagram 1 | Check the status register during power off for low volatge as per datasheet | L | Words should not match |  |  | Uses QPSI. Check voltage via register read during power off. | Firmware test
5 | System | Memory | Read/Write | Functional | Self Test | Check if the memory is able to store the values correctly | Diagram 1 | Enable the memory and write random words to random location; Read back these words | M | Words should match |  |  | Checker board test | Firmware test
6 | System | AD5940 | SPI communication | Functional | Self Test | Read ID | Diagram 1 | Read the AD5940 ID | M | Correct ID is read back |  |  | Already checked from drivers | Final Functional
7 | System | AD5940 | GPIOs | Functional | Self Test | Test the AD5940 GPIOs for register read/write | Diagram 1 | Toggle the GPIOs High and low and check the levels | L | Levels matches the setting |  |  |  | Final Functional
8 | BioZ | DG2502 for AD5940 | Test Switch | Functional | Signal chain integrity test | Turn off AD8233 and ADPD4002 Switches; Turn on and off the 5940 Switch and check if AD5940 is diconneted from the electrodes | Diagram 1 | Use AIN1 and AIN2 to inject a 5 Hz sinusoidal waveform; amp: 0.3Vpp; offset: 1.8V and common mode DC 1.1V sigal via Interface board in AD5940 and check the output | H | Output should match the input when swtich on and show noise when switch off amplitude < -80dB |  |  |  | Final Functional
9 | BioZ | AD5940 and DG2502 | AD5940 accuracy | Performance |  | Enable 5940 Switch and Implement frequncy Sweep and obtain value of a Known Impedance; contact impedance zero Ohm | Diagram 1 | Sweep across 5KHz, 25KHz, 50KHz, 100KHz, 150KHz, 180KHz, 190KHz and 200KHz. Test with 100 Ohm, 510 Ohm and 1kOhm | H | TBD |  |  | For production test use only few points | Characterization
10 | BIOZ | AD5940 and DG2502 | AD5940 accuracy with non-ideal contact impedance | Performance |  | Enable 5940 Switch and Implement sweep in frequency and obtain the value of a known impedance; contact impedance > 0 Ohm | Diagram 1 | Sweep across 5KHz, 25KHz, 50KHz, 100KHz, 150KHz, 180KHz, 190KHz and 200KHz. Test with 100 Ohm, 510 Ohm and 1kOhm | H | TBD |  |  | Not needed for production | Characterization
| BIOZ |  | AD5940 | Performance
11 | BIOZ | AD5940 and DG2502 | AD5940 repeatabiliy | Performance |  | Enable 5940 Switch and Implement sweep in frequency and obtain the value of a known impedance | Diagram 1 | Post process the data from previous measurement and analyze the standard deviation | H | TBD |  |  | Repeat each meausrement 20 or so times for characterization. Repeat only a few times for produciton. | Characterization
| BIOZ |  | AD5940 repeatabiliy | Performance
12 | System | ADXL362 | SPI communication | Functional | Self Test | Read ID | Diagram 1 | Read the ADXL362 ID | L | Correct ID is read back |  |  |  | Final Function
13 | System | ADXL362 | Self Test | Functional | Self Test | Run the ADXL362 self test | Diagram 1 | Run ADXL362 self test | H | Pass criteria as per datasheet |  |  |  | FF/Module
DIAGRAM 1
14 | DG2502 for AD8233 | Test Switch For AD8233 | DG2502 | Functional |  | Turn off AD5940 and ADPD4002 Switches; Turn on and off the 8233 Switch and check if AD8233 is diconneted from the electrodes | Diagram 1 | Close switch. Inject sine wave 5Hz, 1mVpp, Vdc = 1.1V, and check amplitude matches expected ampllitude (+-10%) Open switch. ECG output is lost (no power at 5Hz). Inputs float so there may be energy at 50Hz | H | amplitude < -80dB |  |  | This test should be performed before AD5940 Switch Test | Final Functional
15 | ECG | AD8233 and DG2502 | AD8233 Noise | Performance |  | Enale 8233 switch, connect the contacts to ECG out of interface board, inject know signal value. Collect ADC output data for 61 seconds | Diagram 1 | ECG Differential : Offset - 1.8V, Amplitude- 0.01Vpp, Frequency- 0.1Hz ECG Common: 1.1Vpp, DC; Clip 50 second signal data and Calculate Vpp for the 11 second signal | H | Noise < 50uVpp |  |  | IEC2-27: Input Noise < 30uVpp for at least 10s measurement. Can do shorter test for production? Dan: Yes; not sure how long | Characterization |  | Diagram 1
Performance
16 | ECG | AD8233 and DG2502 | AD8233 Bandwidth | Performance |  | Enable 8233 Switch and connect the contacts to the ECG out of interface board and check the Bandwidth Attenuation from 0.5Hz to 100Hz | Diagram 1 | ECG Differential : Offset - 1.8V, Amplitude- 0.3Vpp, Frequency- 100Hz, 70Hz, 40Hz, 30Hz, 20Hz, 10Hz, 5Hz, 0.5Hz: ECG Common: 1.1Vpp, DC; Vpp at each frequency should match the theortical value | H | Monitor BW 0.34Hz - 38Hz Difference <TBD |  |  | IEC2-27: Frequency and impulse response BW: 0.67Hz to 40Hz with 71% to 110% of output obtained at 5Hz | Characterization
| ECG | AD5940 & AD8233 & DG2502 | AD8233 Bandwidth | Performance
17 | ECG | AD8233 and DG2502 | AD8233 dynamic range | Performance |  | Enable the 8233 Switch and connect the contacts to the ECG out of interface board and check if the AD8233 gain is "constant" | Diagram 1 | Test at 5Hz, 0.5mVpp, 1mVpp & 2mVpp & 5mVpp (5mVpkpk*296=1.48V=+-0.74V) Measure Vrms at 5Hz for each amplitude | H | gain tolerance +-10% |  |  | IEC2-27: Input dynamic range +-10% over specified DC offset | Characterization
18 | ECG | AD8233 and DG2502 | AD8233 Repeatability | Performance |  | Enable the 8233 Switch and connect the contacts to the ECG out of interface board and check if the AD8233 repetability is good | Diagram 1 | Test at 5Hz, 1mVpp . Repeat 10 times. Calculate Std Dev | H | SD < TBD |  |  | IEC2-27: Accuracy of signal reproduction +-20% or +-100uV
19 | ECG | AD8233 and DG2502 | AD8233 CMRR | Performance |  | Enable the 8233 Switch and connect all the contacts to either P or N ECG out of the interface board and measure how good is the AD8233 eliminating the 50-60 Hz noise | Diagram 1 | Drive a 50Hz sinewave Vdc=1.1V Vac=0.3Vpp. Check 20LOG( (amplitude_ecg_codes/gain)/ amplitude_input signal ). Test with no electrode imbalance . Target >80dB - Repeat with 51k¦¦47nF imbalance in one electrode. Target > 40dB | H | CMRR>80dB (no electrode imbalance) CMRR>40dB 51k¦¦47nF electrode imbalance) |  |  | IEC2-27: CMRR
20 | ECG | AD8233 and DG2502 | AD8233 input impedance | Performance |  | Enable the 8233 Switch and connec the contacts to the ECG out of interface board | Diagram 1 | Test at 5Hz, 1mVpp. Check output amplitude. Add 1M in series with each input and check change in output amplitude. Calculate Zin | H | >10M |  |  | IEC2-27: Input Impedance >2.5MoHm
21 | ECG | AD8233 | AD8233 RLD | Performance |  | Check RLD bandwidth and gain with the scope | Diagram 1 | Drive sinewave 10mVpp at different freqeuncies and check volatge at RLD. Comapre to theoretical values. Frequency and Amplitudes to be confirmed. Internal R in AD8233 specifed at +-30% | H | TBD |  |  | Characterize at a couple points | FF/Module
22 | System | ADPD4002 | SPI communication | Functional | Self Test | Read ID | Diagram 1 | Read ADPD4002 ID | L | Correct ID is read back |  |  |  | Final Functional
23 | System | ADPD4002 | GPIOs | Functional | Self Test | Test the ADPD4002 GPIOs are soldered correctly | Diagram 1 | Toggle the GPIOs High and low and check the levels | L | Levels matches the setting |  |  |  | Final Functional
24 | System | ADPD4002 | ADC Vref | Functional | Self Test | Use Vc2 connected at ET1/EB2 and measure the value | Diagram 1 | TBD | M | as per datasheet |  |  |  | Characterization
25 | System | ADPD4002 and LEDs | Chcek the led Green and Red LED on/off | Functional | User Interface | Toggle the LEDs on and off | Diagram 1 | Toggle LEDs on/off and make a visual observation if they are as per setting | H | LEDs turn On and Off |  |  |  | Final Functional
26 | System | ADPD4002 and LEDs | Check the IR LED on/off | Fucntional | User Interface | Use a black box to cover the LEDS and turn on the IR led collect the reflection on a PD connected to the black box | Diagram 1 | Pass the known excitation to the IR LED and oberve the equivalent output at the PD | H | PD output should match calculated value |  |  |  | Final Functional
27 | System | ADPD4002 and PD | Check the dark current and light on the PD | Functional | User Interface | Use a black box to cover the LEDS and turn on/off the green LED connected to the box; collect the dark current and light reflection on a PD | Diagram 1 | Pass the known exciation current through the green LED on the box and measure the dark current and light reflection | H | PD output should match the datasheet |  |  |  | Final Functional
28 | DG2502 for ADPD4002 | Test Switch for ADPD4002 | DG2502 | Functional |  | Turn off AD8233 and AD5940 Switches; Turn on and off the 4002 Switch and check if ADPD4002 is diconneted from the electrodes | Diagram 1 | Close switch. Inject sine wave 5Hz, 1mVpp, Vdc = 1.1V, and check amplitude matches expected ampllitude (+-10%) Open switch. ECG output is lost (no power at 5Hz). Inputs float so there may be energy at 50Hz | M | amplitude < -80dB |  |  |  | Final Functional
29 | BioZ | ADPD4002 and DG2502 | Accuracy | Performance |  | Enable 4002 Switch and Implement frequncy Sweep and obtain value of a Known Impedance; contact impedance zero Ohm | Diagram 1 | Sweep across 5KHz, 25KHz, 50KHz, 100KHz, 150KHz, 180KHz, 190KHz and 200KHz. Test with 100 Ohm, 510 Ohm and 1kOhm | M | TBD |  |  |  | Characterization/FF- only few points
30 | BIOZ | ADPD4002 and DG2502 | ADPD4002 accuracy with non-ideal contact impedance | Performance |  | Enable 5940 Switch and Implement sweep in frequency and obtain the value of a known impedance; contact impedance > 0 Ohm | Diagram 1 | Sweep across 5KHz, 25KHz, 50KHz, 100KHz, 150KHz, 180KHz, 190KHz and 200KHz. Test with 100 Ohm, 510 Ohm and 1kOhm | M | TBD |  |  |  | Characterization/FF- only few points
| BIOZ |  | AD5940 | Performance
31 | BIOZ | ADPD4002 and DG2502 | ADPD4002 repeatabiliy | Performance |  | Enable 4002 Switch and Implement sweep in frequency and obtain the value of a known impedance | Diagram 1 | Post process the data from previous measurement and analyze the standard deviation | M | TBD |  |  |  | Characterization/FF- only few points
| BIOZ |  | AD5940 repeatabiliy | Performance
32 | ECG | ADPD4002 and DG2502 | ADPD4002 Noise | Performance |  | Enale 4002 switch and using interface board produce know value of differential and common mode signal. Collect ADC output data for 61 seconds | Diagram 1 | ECG Differential : Offset - 1.8V, Amplitude- 0.01Vpp, Frequency- 0.1Hz ECG Common: 1.1Vpp, DC; Clip 50 second signal data and Calculate Vpp for the 11 second signal | M | Noise < 50uVpp |  |  |  | Characterization
Performance
33 | ECG | ADPD4002 and DG2502 | ADPD4002 Bandwidth | Performance |  | Enable 4002 Switch and connect the contacts to the ECG out of interface board and check the Bandwidth Attenuation from (same)?? | Diagram 1 | ECG Differential : Offset - 1.8V, Amplitude- 0.3Vpp, Frequency- 40Hz, 30Hz, 20Hz, 10Hz, 5Hz, 0.5Hz, 0.1Hz : ECG Common: 1.1Vpp, DC; Vpp at each frequency should match the theortical value | M | Monitor BW 0.34Hz - 38Hz Difference <TBD
| ECG | AD5940 & AD8233 & DG2502 | AD8233 Bandwidth | Performance
34 | ECG | ADPD4002 and DG2502 | ADPD4002 Dynamic range | Performance |  | Enable the 4002 Switch and connect the contacts to the ECG out of interface board and check if the ADPD4002 gain is "constant" | Diagram 1 | Test at 5Hz, 0.5mVpp, 1mVpp & 2mVpp & 5mVpp (5mVpkpk*296=1.48V=+-0.74V) Measure Vrms at 5Hz for each amplitude | M | gain tolerance ??
|  | AD5940 & AD8233 & DG2502 | AD8233 Bandwidth | Performance
35 | ECG | ADPD4002 and DG2502 | ADPD4002 Repeatability | Performance |  | Enable the 4002 Switch and connect the contacts to the ECG out of interface board and check if the ADPD4002 repetability is good | Diagram 1 | Test at 5Hz, 1mVpp . Repeat 10 times. Calculate Std Dev | M | SD < TBD
| ECG | AD8233 and DG2502 | AD8233 Repeatability | Performance |  | Enable the 8233 Switch and connect the contacts to the ECG out of interface board and check if the AD8233 repetability is good |  | Test at 5Hz, 1mVpp . Repeat 10 times. Calculate Std Dev |  | SD < TBD
36 | ECG | ADPD4002 and DG2502 | ADPD4002CMRR | Performance |  | Enable the4002 Switch and connect all the contacts to either P or N ECG out of the interface board and measure how good is the ADPD4002 eliminating the 50-60 Hz noise | Diagram 1 | Drive a 50Hz sinewave Vdc=1.1V Vac=0.3Vpp. Check 20LOG( (amplitude_ecg_codes/gain)/ amplitude_input signal ). Test with no electrode imbalance . Target >80dB - Repeat with 51k¦¦47nF imbalance in one electrode. Target > 40dB | M | CMRR>80dB (no electrode imbalance) CMRR>40dB 51k¦¦47nF electrode imbalance)
37 | ECG | ADPD402 and DG2502 | ADPD4002 input impedance | Performance |  | Enable the 4002 Switch and connec the contacts to the ECG out of interface board | Diagram 1 | Test at 5Hz, 1mVpp. Check output amplitude. Add 1M in series with each input and check change in output amplitude. Calculate Zin | M | >10M
38 | System | DIALOG | Test Bluetooth link | Functional | user feedback | RF section is functional | Diagram 1 | reset dialog, load program (advertising). Check RF link works with USB dongle /phone or other | L | system advertises |  |  | This test needs some examinaion if it could be automated; test signal strength? | Final Functional
|  | DIALOG | Test Bluetooth link | Functional | user feedback | RF section is functional |  | reset dialog, load program (advertising). Check RF link works with USB dobngle /phone or other |  | system advertises
39 | System | DIALOG | UART/SPI interface | Functional | Self Test | Uart loopback test | Diagram 1 | reset dialog, load program (uart loopback). SPI and UART works | L | program downloads and words match |  |  |  | Final Functional
| System | DIALOG | Test Bluetooth link | Functional | user feedback | RF section is functional |  | reset dialog, load program (uart loopback). SPI and UART works |  | system advertises
40 | System | Button | Press button | Functional | user feedback | Push buttons and check response | Diagram 1 | ask user to push button and wait for interrupt | L | interrupt received |  |  |  | Final Functional
41 | Temperature | TEMP. SENSING | Temp Sensing linearity alone | Performance |  | Check circuit linearity simulating the NTC with a Programmable Resistance substitutter or Zbank | Diagram 1 | change the resistor value and obtain the temperature value. Check if the adquired values matches with the expected one | H | tbd |  |  | Use Coolidge for temperature testing?
42 | Temperature | TEMP. SENSING | Temp Sensing Repeatability alone | Performance |  | Check repeatibility simulating the NTC with a Programmable Resistance substitutter or Zbank | Diagram 1 | measure several times the same value and obtain the std deviation | H | tbd |  |  |  | Characterization
43 | Temperature | TEMP. SENSING | Temp Sensing linearity all togheter | Performance |  | Check circuit linearity simulating the NTC with a Programmable Resistance substitutter or Zbanck while the ECG and Z signals are also adquired. Is the noise floor the same? | Diagram 1 | change the resistor value and obtain the value. Check if the adquired values matches with the expected one | L | tbd |  |  |  | Characterization
44 | Temperature | TEMP. SENSING | Temp Sensing Repeatability all togheter | Performance |  | Check repeatibility simulating the NTC with a Programmable Resistance substitutter or Zbanck while the ECG and Z signals are also adquired. Is the noise floor the same? | Diagram 1 | measure several times the same value and obtain the std deviation | L | tbd |  |  |  | Characterization/FF- only few points
45 | AD7156 | I2C communication | Read ID | Functional | Self Test | Read device ID | Diagram 1 | Read the device ID | L | ID is correct |  |  |  | Final Functional
46 | AD7156 | Value out | Use a metal plate to generate an output signal |  |  | Check the output with and without a metal plate placed below the watch | Diagram 1 | TBD | L | Output is observed |  |  |  | Module
48 | ADP5360 | I2C communication | Read ID | Functional | Self Test | Read device ID | Diagram 1 | Read the device ID | L | ID is correct |  |  |  | Final Functional
49 | POWER MANAGEMENT | POWER MANAGEMENT & USB POWER | USB "wall" and "PC" charges battery test | Functional |  | check the USB "Wall" and "PC" supply recharges the battery (USB "PC" and "Wall" connected alternately) | Diagram 1 | mimick battery with keithley and check it sinks current (so USB5V is present and charging the battery) | M | charging current >80mA |  |  | follow up with Dan | Final Functional
50 | POWER MANAGEMENT | POWER MANAGEMENT & USB POWER | USB PC test | Functional |  | check the USB "PC" connection (battery is connected, USB "PC" is connected but USB "Wall" is not connected) | Diagram 1 | set keithley ch1 to 3.8 volts, keithley ch2 is disconnected and ch3 is connected (4.8v) and OTG_SEL means not get voltage from PC. check ch3 current is zero or negative and check USB_ID. Second step: change OTG_SEL, therefore ch3 current will have to be positive, USB_ID will change its level and the USB communication has to work. | M | ADP5350 status is correct |  |  |  | Final Functional
51 | POWER MANAGEMENT | POWER MANAGEMENT & USB POWER | USB Wall test | Functional |  | check the USB "Wall" connection (battery is connected, USB "PC" not connected but USB "Wall" is connected) | Diagram 1 | set keithley ch1 to 3.8 volts, keithley ch2 is connected (5volts) and ch3 is connected (4.8v) and OTG_SEL low. check ch3 current is zero or negative and check USB_ID is floating (set input with pull-up). Second step: change OTG_SEL to high. every thing remains equal. Third step: Ch2 is disconnected, USB_ID goes low and ch3 will supply the FTDI. send message and wait response by USB. | M | ADP5350 status is correct |  |  | 25mV resolution of input. Can possibly utilize AD8233 | Final Functional
52 | POWER MANAGEMENT | POWER MANAGEMENT & USB POWER | Battery test | Functional |  | check supplies. Just the battery is connected to the board (USB "Wall" not connected and USB "PC" not connected) | Diagram 1 | Check if the watch is working without connecting to the Wall or PC charger | L | ADP5350 status is correct |  |  |  | Final Functional
53 | POWER MANAGEMENT |  | Power level test |  |  | Test the power drawn by the board at different power modes |  | TBD | L | TBD |  |  |  | Final Functional
54 | ADP160 | Relibility | Test that LDO produces required volatge when input is changed | Performance |  | Simply toggle input to make sure output is constant | Diagram 1 | Not only can we vary Vin while checking Vout, we can also test the LDO with and without a load. Load a 1k and 100k resistor to Vout to make sure Vout remains the same. |  | equal to Vout ± 5% |  |  |  | Final Functional
Electrodes

## Sheet3
101.5:Crosstalk
101.6: Gain Control and Stability
101.7: Sweep Speed
101.9: Gain Indicator
101.11: Baseline Reset
101.12: Pacemaker pulse display capability

## Sheet2
r4 | is | not
r5 | is | not
c12 | not | is
c10 | not | is
c5 | not | is

## Performance
Functions | Test | Watch PCB controls | DAC0 | DAC1 | GPIOs | Procedure | Pass criteria | Pass/Fail | Comments
ECG | AD8233 Switch functionality | Turn off AD5940 and ADPD4002 Switches; Turn on and off the 8233 Switch and check if AD8233 is disconneted from the electrodes | Sinwave(pos) freq:5Hz, amp= 0.1Vpp | Sinwave(neg) freq:5Hz, amp= 0.1Vpp | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Close DG2502 switch and collect data for 10 sec Open DG2502 switch and collect data for 10 sec | Close switch :amplitude matches expected ampllitude (+-10%) Open switch: ECG output is lost (no power at 5Hz). Inputs float so there may be energy at 50Hz |  | This test should be performed before AD5940 Switch Test
| ECG Signal quality with ECG simulator | Turn off AD5940 and ADPD4002 Switches; Turn on the 8233 Switch and collect AD5940 output data | - | - | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:1-2 JP8:2-3 | Display the waveform and record the data for further examination | Take FFT should not see any 50,60Hz noise
| ECG Signal quality on body | Turn off AD5940 and ADPD4002 Switches; Turn on the 8233 Switch and collect AD5940 output data | - | - | SNOISE1:0 SNOISE2: 0 | Connect one end wet electrodes to Pins 8,6 and 5 of JP14 and place the wet electrodes on left/right wrist and two fingers of the other hand. Display the waveform and save the data for further examination | Match the data with the best waveform recorded by Gen 3 watch
Accelerometer | Self Test for ADXL362 | Run the self test for ADXL362 | - | - | - | Run the ADXL362 self test | Pass criteria as per datasheet
PPG | Dark Test | Turn on ADPD to measure PPG | - | - | - | Place a refective surface over the LEDs and measure the output at 1 pulse 0mA current | dBm >= 85
| Test without AGC | Turn on ADPD to measure PPG | - | - | - | Place the finger/wrist on the optical window and display PPG using wavetool | Good signal (close to finger clip) On Praptee, Shane, Supriya, Dan, Foroohar
| Test with AGC | Turn on ADPD to measure PPG | - | - | - | Place the finger/wrist on the optical window and display PPG using wavetool | Good signal (close to finger clip) On Praptee, Shane, Supriya, Dan, Foroohar
Temperature | Accuracy | Place a metal bar with (palladin) over the RTD and change the temperature | - | - | -
BLE | Functionality | Connect BLE to Watch PCB and collect data |  |  |  | Collect data for an hour of ECG, PPG, Temperature and ADXL at all the frequencies to be demod at CES | BLE stays connected and no packet lose reported
| Reliability | Collect data over BLE from N channels (as recommended by Firmware team) over night | - | - | - | Collect PPG, Temperature and ADXL data for an hour and ECG data for another hr | No data lost
Battery Testing |  | Follow Fenda's tests

## ECG
Test | Watch PCB Controls | DAC0 | DAC1 | GPIO | Procedure | Priority | Pass Criteria | Pass/Fail | Comments
DG2502 | Turn off AD5940 and ADPD4002 Switches; Turn on and off the 8233 Switch and check if AD8233 is diconneted from the electrodes | Sinwave(pos) freq:5Hz, amp= 0.1Vpp | Sinwave(neg) freq:5Hz, amp= 0.1Vpp | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Close DG2502 switch and collect data for 10 sec Open DG2502 switch and collect data for 10 sec | H | Close switch :amplitude matches expected ampllitude (+-10%) Open switch: ECG output is lost (no power at 5Hz). Inputs float so there may be energy at 50Hz |  | This test should be performed before AD5940 Switch Test amplitude < -80dB
AD8233 Noise | Enale 8233 switch, connect the contacts to ECG out of interface board, inject know signal value. Collect ADC output data for 61 seconds | NA | NA | SNOISE1:0 SNOISE2: 1 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Connect all the electrodes and RLD will provide 1.1Vdc; collect 50 second signal data and Calculate Vpp for the 11 second signal | H | Noise < 50uVpp |  | IEC2-27: Input Noise < 30uVpp for at least 10s measurement. Can do shorter test for production? Dan: Yes; not sure how long
AD8233 Bandwidth | Enable 8233 Switch and connect the contacts to the ECG out of interface board and check the Bandwidth Attenuation from 0.5Hz to 40Hz | Sinwave(pos) freq:40Hz - 0.5Hz, amp= 0.3Vpp | Sinwave(neg) freq:40Hz - 0.5Hz, amp= 0.3Vpp | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Test Frequencies- 40Hz, 30Hz, 20Hz, 10Hz, 5Hz, 0.5Hz: ECG Common: 1.1Vpp, DC; Vpp at each frequency should match the theortical value | H | Monitor BW 0.34Hz - 38Hz Difference <TBD |  | IEC2-27: Frequency and impulse response BW: 0.67Hz to 40Hz with 71% to 110% of output obtained at 5Hz
AD8233 Bandwidth
AD8233 dynamic range | Enable the 8233 Switch and connect the contacts to the ECG out of interface board and check if the AD8233 gain is "constant" | Sinwave(pos) freq:5Hz, amp= 0.1Vpp-0.5Vpp | Sinwave(neg) freq:5Hz, amp= 0.1Vpp-0.5Vpp | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Test amplitudes: 0.5Vpp, 0.1Vpp & 0.2Vpp & 0.5Vpp Measure Vrms at 5Hz for each amplitude | H | gain tolerance +-10% |  | IEC2-27: Input dynamic range +-10% over specified DC offset
AD8233 Repeatability | Enable the 8233 Switch and connect the contacts to the ECG out of interface board and check if the AD8233 repetability is good | Sinwave(pos) freq:5Hz, amp= 0.1Vpp | Sinwave(neg) freq:5Hz, amp= 0.1Vpp | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Test at 5Hz, 1mVpp . Repeat 10 times. Calculate Std | H | SD < TBD |  | IEC2-27: Accuracy of signal reproduction +-20% or +-100uV
AD8233 CMRR | Enable the 8233 Switch and connect the contacts to ECG out of the interface board and measure how good is the AD8233 eliminating the 50-60 Hz noise | Sinwave(pos) freq:50Hz, amp= 0.1Vpp | Sinwave(pos) freq:50Hz, amp= 0.1Vpp | testa: SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 testb: SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 0 ECG_POSSEL: 1 testc: SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 0 JP7:2-3 JP8:1-2 | Check 20LOG( (amplitude_ecg_codes/gain)/ amplitude_input signal ). Test with no electrode imbalance . Target >80dB - Repeat with 51k¦¦47nF imbalance in one electrode. Target > 40dB | H | CMRR>80dB (no electrode imbalance) CMRR>40dB 51k¦¦47nF electrode imbalance) |  | IEC2-27: CMRR
AD8233 input impedance | Enable the 8233 Switch and connec the contacts to the ECG out of interface board | Sinwave(pos) freq:5Hz, amp= 0.1Vpp | Sinwave(neg) freq:5Hz, amp= 0.1Vpp | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Replace R25 and R26 by 1M resistor. Check output amplitude. Add 1M in series with each input and check change in output amplitude. Calculate Zin | H | >10M |  | IEC2-27: Input Impedance >2.5MoHm

## ADPD
Test | Watch PCB Controls | DAC0 | DAC1 | GPIO | Procedure | Priority | Pass Criteria | Pass/Fail | Comments
DG2502 for ADPD4002 | Test Switch for ADPD4002 | Sinwave(pos) freq:5Hz, amp= 0.1Vpp | Sinwave(neg) freq:5Hz, amp= 0.1Vpp | SNOISE1:0 SNOISE2: 0 ECG_NEGSEL: 0 ECG_POSSEL: 0 JP7:2-3 JP8:1-2 | Close DG2502 switch and collect data for 10 sec Open DG2502 switch and collect data for 10 sec | H | Close switch :amplitude matches expected ampllitude (+-10%) Open switch: ECG output is lost (no power at 5Hz). Inputs float so there may be energy at 50Hz
Read device ID | Read the ADPD ID | NA | NA | SNOISE1:0 SNOISE2: 0 ECG_NEGSEL: 0 ECG_POSSEL: 0 JP7:2-3 JP8:1-2 | Read ADPD4002 ID | H | Correct ID is read back

## PPG
Test | Watch PCB Controls | DAC0 | DAC1 | GPIO | Procedure | Priority | Pass Criteria | Pass/Fail | Comments
Dark current | Set all current values to 0mA; frequency to 100Hz and No. of pulses to 1; turn on both the channels | NA | NA | SNOISE1:0 SNOISE2: 0 ECG_NEGSEL: 0 ECG_POSSEL: 0 JP7:2-3 JP8:1-2 | Collect the data for 10sec and check the standard deviation for both the channels | H | std < 2
LED Drivers | Place the finger clip board on the HR board at some distance and turn LEDs one at a time set the current to 2mA, number of pulses 16 | NA | NA | SNOISE1:0 SNOISE2: 0 ECG_NEGSEL: 0 ECG_POSSEL: 0 JP7:2-3 JP8:1-2 | Collect the data for 10sec on the finger clip board PD and match the current value | H | Match the value of the PD output to the input current
PD | Place the finger clip board on the HR board at some distance and turn on green LED on the finger clip board to 2mA, 16 pulses at a time set the current to 2mA, number of pulses 16 and turn on one channel at a time on the HR board | NA | NA | SNOISE1:0 SNOISE2: 0 ECG_NEGSEL: 0 ECG_POSSEL: 0 JP7:2-3 JP8:1-2 | Collect the data for 10sec on HR board and match the input current to PD output | H | Match the value of the PD output to the input current

## BioZ
Test | Watch PCB Controls | DAC0 | DAC1 | GPIO | Procedure | Priority | Pass Criteria | Pass/Fail | Comments
Test Switch | Turn off AD8233 and ADPD4002 Switches; Turn on and off the 5940 Switch and check if AD5940 is diconneted from the electrodes | Sinwave(pos) freq:5Hz, amp= 0.1Vpp | Sinwave(neg) freq:5Hz, amp= 0.1Vpp | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Close DG2502 switch and collect data for 10 sec Open DG2502 switch and collect data for 10 sec | H | Output should match the input when swtich on and show noise when switch off amplitude < -80dB
AD5940 accuracy | Enable 5940 Switch and Implement frequncy Sweep and obtain value of a Known Impedance; contact impedance zero Ohm | Sinwave(pos) amp= 0.1Vpp; Sweep frequency | Sinwave(neg) amp= 0.1Vpp; Sweep frequency | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Sweep across 5KHz, 25KHz, 50KHz, 100KHz, 150KHz, 180KHz, 190KHz and 200KHz. Test with 50 Ohm, 100 Ohm and 1kOhm | H | Magnitude accuracy within 3% Phase accuracy within 6 degrees
AD5940 repeatabiliy | Enable 5940 Switch and Implement frequncy Sweep and obtain value of a Known Impedance; contact impedance 0 Ohm | Sinwave(pos) amp= 0.1Vpp; Sweep frequency | Sinwave(neg) amp= 0.1Vpp; Sweep frequency | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Sweep across 5KHz, 25KHz, 50KHz, 100KHz, 150KHz, 180KHz, 190KHz and 200KHz. Test with 50 Ohm, 100 Ohm and 1kOhm; repeat each reading 10 times | H | Magnitude standard deviation < 0.1Ohm and Phase standard variation <0.5 degree
AD5940 repeatabiliy
AD5940 accuracy with non-ideal contact impedance | Enable 5940 Switch and Implement frequncy Sweep and obtain value of a Known Impedance; contact impedance 1kOhm, 2kOhm and 3kOhm | Sinwave(pos) amp= 0.1Vpp; Sweep frequency | Sinwave(neg) amp= 0.1Vpp; Sweep frequency | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Sweep across 5KHz, 25KHz, 50KHz, 100KHz, 150KHz, 180KHz, 190KHz and 200KHz. Test with 100 Ohm, 50 Ohm and 1kOhm | H | Magnitude accuracy within 3% Phase accuracy within 6 degrees |  | Pass criteria might change
AD5940
AD5940 accuracy with contact impedance mismatch | Enable 5940 Switch and Implement frequncy Sweep and obtain value of a Known Impedance; contact impedance set F+/S+/F-/S- to 2/2/2/1 kOhm | Sinwave(pos) amp= 0.1Vpp; Sweep frequency | Sinwave(neg) amp= 0.1Vpp; Sweep frequency | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Sweep across 5KHz, 25KHz, 50KHz, 100KHz, 150KHz, 180KHz, 190KHz and 200KHz. Test with 100 Ohm, 50 Ohm and 1kOhm | H | Magnitude accuracy within 3% Phase accuracy within 6 degrees (pass criteria might change) |  | Change the mismatch impedance to all the four contacts and repeat the test
AD5940

## Temperature
Test | Watch PCB Controls | DAC0 | DAC1 | GPIO | Procedure | Priority | Pass Criteria | Pass/Fail | Comments
Accuracy /Reliability | Place the Palladin above the temp sensor increase and decrease the temperature in steps and match the value from ADPD output | NA | NA | NA |  | H | Temperature value should match within 1 degree

## Battery
Test | Watch PCB Controls | DAC0 | DAC1 | GPIO | Procedure | Priority | Pass Criteria | Pass/Fail | Comments
Charging Current Consumption | enter PCB into specified mode. | NA | NA | NA | Connect Keithley (V/ Source/Meter) to PCB |  | Verify that Charging Current Limits [-45mA:-14mA]
Static Current Consumption | enter PCB into specified mode. | NA | NA | NA | Connect Keithley (V/ Source/Meter) to PCB |  | Verify that Static Current Limits [1mA:5mA]
Hibernate Current Consumption | enter PCB into specified mode. | NA | NA | NA | Connect Keithley (V/ Source/Meter) to PCB |  | Verify that Hibernate Current Limits [1mA:5mA]
Sweep Stage Current Consumption | enter PCB into specified mode. | NA | NA | NA | Connect Keithley (V/ Source/Meter) to PCB |  | Verify Current Limits
Streaming Stage Current Consumption | enter PCB into specified mode. | NA | NA | NA | Connect Keithley (V/ Source/Meter) to PCB |  | Verify Current Limits
BLE advertising Stage Current Consumption | enter PCB into specified mode. | NA | NA | NA | Connect Keithley (V/ Source/Meter) to PCB |  | Verify Current Limits

## Schedule
Sr.No# |  |  | Comments | Status | Resourses
| Performance Tests | 2020-01-23 |  | . | Nitin, Praptee and Shane; Dan as needed | Open
1 | AD8233 Switch functionality |  |  | Implemented |  | In Progress
2 | ECG Signal quality with ECG simulator |  |  | Implemented |  | Implemented
3 | ECG Signal quality on body |  |  | Implemented |  | Tested
4 | Self Test for ADXL362 |  | Will be added by Wednesday | Implemented |  | Completed
5 | PPG Dark Test |  |  | Implemented
6 | PPG Test without AGC |  |  | Implemented
7 | PPG Test with AGC |  | depends on Firmware | Implemented
8 | Temperature Accuracy |  | will be done as part of temperature testing | Implemented
9 | BLE Functionality | 2020-01-31 | Need to discuss test parameters with the team
10 | BLE Reliability
11 | Battery
| ECG for AD8233 | 2020-01-31 |  | .
12 | 8233 switch (DG2502) |  | implemented in performance Tests | Implemented
13 | AD8233 input Noise |  | Will be ready by Monday | Implemented
14 | AD8233 Bandwidth |  |  | Implemented
15 | AD8233 dynamic range |  |  | Implemented
16 | AD8233 Repeatability |  |  | Implemented
17 | AD8233 CMRR |  |  | Implemented
18 | AD8233 input impedance |  |  | Implemented
| BioZ for AD5940 | 2-3days | depends on Firmware to release the impedance test app | .
19 | 5940 switch (DG2502) |  |  | Implemented
20 | AD5940 accuracy |  |  | Implemented
21 | AD5940 repeatabiliy |  |  | Implemented
22 | AD5940 accuracy with non-ideal contact impedance |  |  | Implemented
23 | AD5940 accuracy with contact impedance mismatch |  |  | Implemented
| Battery | 2020-02-05 | Will be ready by Wednesday | .
24 | Charging Current Consumption
25 | Static Current Consumption
26 | Hibernate Current Consumption
27 | Sweep Stage Current Consumption
28 | Streaming Stage Current Consumption
29 | BLE advertising Stage Current Consumption
| Temperature | 2020-01-31 |  | .
30 | Accuracy /Reliability |  |  | Implemented
| PPG | 2-3days | do we need these tests ?? | .
31 | Dark current
32 | LED Drivers
33 | PD
| ECG for ADPD4000 | 2-3days | depends on Firmware to release the ADPD BioZ test app | .
34 | 8233 switch (DG2502)
35 | AD8233 input Noise
36 | AD8233 Bandwidth
37 | AD8233 dynamic range
38 | AD8233 Repeatability
39 | AD8233 CMRR
40 | AD8233 input impedance
| BioZ for ADPD4000 | 2-3days | depends on Firmware to release the ADPD ECG test app | .
41 | 5940 switch (DG2502)
42 | AD5940 accuracy
43 | AD5940 repeatabiliy
44 | AD5940 accuracy with non-ideal contact impedance
45 | AD5940 accuracy with contact impedance mismatch