## Test_list_updated
Sr.No# | Test | WATCH | PCB | PPG | PPG + Bio-Z | PPG + ECG | ECG | Test methodology | Aspects to be Analyzed | Pass/Fail creteria | Status | Time for completion | Resource | Priority | Test type | Start | End | Comments
ECG Tests
1 | ECG Signal quality with ECG simulator | All | All |  |  |  |  | check 'Performance' sheet of this excel for details | a. Presence of the correct ECG morphology; b. Strength of P,QRS and T waves; c. ECG does not clip or rail | a. P wave seen? QRS wave seen? T wave seen?; b. T_SNR > 3dB R_SNR > 20dB and 0.3 < SDR < 0.95 c.Valid ECG data > 80% of each recording for all recordings | Paritally completed. Data can be parsed manually after test. Working on Automation. Need to write matlab script for analysis | 2 days | Praptee/ Nitin to implement in automated tests | H | Performance |  |  | - LabVIEW ECG extractor created
2 | ECG Signal quality with ECG simulator in relation with PPG, accelerometer and temperature (use case #4) | All | All | 50Hz |  |  | 1000Hz | Run ECG@1000Hz, PPG@ 50Hz, ADXL@50Hz and temp@1Hz | 1. ECG morphology; 2. ECG SNR; 3. PPG morphology; 4. PPG SNR (MI); 5. Data integrity | 1. ECG morphology is as per requirement; 2. ECG SNR is not reduced; 3. PPG morphology is as per requirement; 4. PPG MI is not reduced; 5. No packet loss | Paritally completed. Data can be parsed manually after test. Working on Automation. Need to write matlab script for analysis | 1 day | Praptee | H | System/Functional
3 | ECG Signal at different frequencies | All | All |  |  |  |  | Check the ECG morphology at 250Hz, 500Hz and 1000Hz; Collect data for a 60 secs at each frequency there | ECG signal quality at the use case frequencies | Correct ECG morphology should be observed and no data loss observed | Not yet implemented | 1 day | Praptee | M | Performance |  | 2020-05-18
4 | ECG Signal quality on body | Few | Few |  |  |  |  | Please check 'Performance' sheet of this excel for details | a. Presence of the correct ECG morphology; b. Strength of P,QRS and T waves; c. ECG does not clip or rail | a. P wave seen? QRS wave seen? T wave seen?; b. T_SNR > 3dB R_SNR > 20dB and 0.3 < SDR < 0.95 c.Valid ECG data > 80% of each recording for all recordings | Paritally completed. Data can be parsed manually after test. Working on Automation. Need to write matlab script for analysis | NA | NA | M | Performance
5 | AD8233 switch (DG2502) | All | All |  |  |  |  | Check 'ECG' tab for details | Is AD8233 switch working | Close switch :amplitude matches expected ampllitude (+-10%) Open switch: ECG output is lost (no power at 5Hz). Inputs float so there may be energy at 50Hz | Done | NA | NA | H | Functional
6 | AD8233 input Noise | All | All |  |  |  |  | Check 'ECG' tab for details | check if the noise floor is within limit | Noise < 50uVpp | Done | NA | NA | H | Functional/Characterization
7 | AD8233 Bandwidth | None | All |  |  |  |  | Check 'ECG' tab for details | Ad8233 bandwidth | Bandwidth attenuation of AD8233 is as per circuit design | Done | NA | NA | H | Charaterization
8 | AD8233 dynamic range | None | All |  |  |  |  | Check 'ECG' tab for details | AD8233 gain is as per circuit design 150+-10% | AD8233 gain is within tolerance | Done | NA | NA | H | Charaterization
9 | AD8233 Repeatability | None | All |  |  |  |  | Check 'ECG' tab for details | Output of ECG circuit (Ad8233+Ad5940) is repetable | SD < TBD | Done | NA | NA | H | Charaterization
10 | AD8233 CMRR | None | All |  |  |  |  | Check 'ECG' tab for details | AD8233's ability to reject 50 and 60Hz noise | CMRR>80dB (no electrode imbalance); CMRR>40dB 51k¦¦47nF electrode imbalance) | Done | NA | NA | H | Charaterization
11 | AD8233 input impedance | None | All |  |  |  |  | Check 'ECG' tab for details | Input impedance | Zin > 10M | Done | NA | NA | H | Charaterization
12 | ADPD4000 switch (DG2502) | All | All |  |  |  |  | Turn off AD5940 and AD8233 Switches; Turn on and off the ADPD4000 Switch and check if it is diconneted from the electrodes | Is ADPD4000 switch working | Close switch :amplitude matches expected ampllitude (+-10%) Open switch: ECG output is lost (no power at 5Hz). Inputs float so there may be energy at 50Hz | Yet to be implemented |  | Praptee & Nitin | L | Functional
13 | ADPD4000 input noise test | All | All |  |  |  |  | Enale ADPD4000 switch, connect the contacts to ECG out of interface board, inject know signal value. Collect ADC output data for 61 seconds | check if the noise floor is within limit | Noise <xuVpp | Yet to be implemented |  | Praptee & Nitin | L | Characterization
14 | Output accuracy (linearity) | All | All |  |  |  |  | Input signals in the range of ±5 mV, varying at a rate up to 125 mV/s, Tests with ±5/2.5/1/0.5 Vpp triangular wave (2/4/6Hz). | Accuracy of Signal reproduction | input should be reproduced on the output ≤ ±20% (≤ ±10% for IEC60601-2-47) of the nominal value of the output or ±100µV (±50µV for IEC60601-2-47), whichever is greater. | Yet to be implemented |  | Praptee & Nitin | L | Characterization
15 | ADPD4000 Frequency response | All | All |  |  |  |  | Enable ADPD4000 Switch and connect the contacts to the ECG out of interface board and check the Bandwidth Attenuation from 0.5Hz to 40Hz | Frequency response of ADPD4000 | the output at 0.34 Hz and 38 Hz shall be within 71% to 110% (±3 dB) of the output obtained with a 5 Hz sine wave input signal | Yet to be implemented |  | Praptee & Nitin | L | Characterization
ADXL Tests
16 | Self Test for ADXL362 | All | All |  |  |  |  | Run the self test | Accelerometer is performing as expected | Test output is PASS | Done |  | Nitin | M | Functional |  | 2020-05-18
PPG Tests
17 | PPG Dark Test | All | All |  |  |  |  | check 'PPG' sheet of this excel for details | Standard deviation | std <2 | Done |  | Nitin | H | Charaterization
18 | PPG Moduation Index | All | All |  |  |  |  | Use data from PPG test with AGC and calculate MI | Ratio of AC part of the PPG signal to the DC part |  | Auto data analyis needs to be implemented |  | Praptee and Nitin | H | Charaterization | Waiting for AGC. Basic test template is created on top of which each of the test cases can be built based on the AGC implementation and functions.
19 | PPG Test with AGC | All | All |  |  |  |  | Using AGC take PPG data on wrist | Check for the correct PPG morphology |  | Needs to be implemented | 3 days | Praptee and Nitin | H | Performance
20 | PPG Test with AGC at 100Hz and 500Hz | All | All |  |  |  |  | Using AGC and take PPG data on wrist | Check for the correct PPG morphology | MI as per requirement | Needs to be implemented
21 | PPG Test with AGC in relation to adxl and temp (use case #1) | All | All | 500Hz |  |  |  | Run PPG@ 500Hz, ADXL@50Hz and temp@1Hz | 1. PPG morphology; 2. PPG SNR; 3. Data integrity | 1. PPG morphology as per requirement; 2. PPG MI does not decrease; 3. no data loss | Needs to be implemented | 1 day Nitin | Praptee and Nitin | H | System/Functional
22 | PPG Test with AGC in relation to EDA, adxl and temp (use case #2) | All | All | 100Hz |  |  |  | Run EDA@30Hz, PPG@ 100Hz, ADXL@50Hz and temp@1Hz | 1. PPG morphology; 2. EDA value is correct;3 PPG SNR (MI); 4. Data integrity | 1. PPG morphology as per requirement; 2. EDA value within range; 3. PPG MI does not decrease; 4. no data loss | Needs to be implemented | 1 day Nitin | Praptee and Nitin | H | System/Functional
23 | PPG Test with AGC in relation to ECG, adxl and temp (use case #3) | All | All | 100Hz |  |  | 250Hz | Run ECG@250Hz, PPG@ 100Hz, ADXL@50Hz and temp@1Hz | 1. PPG morphology; 2. ECG morphology;3 PPG SNR (MI); ECG SNR; 5. Data integrity | 1. PPG morphology as per requirement; 2. ECG morphology as per requirement; 3. PPG MI does not decrease; 4. ECG SNR does not decrease; 5. no data loss | Needs to be implemented | 1 days Nitin | Praptee and Nitin | H | System/Functional
24 | PPG without AGC at 100Hz and 500Hz | All | All |  |  |  |  | Turn off AGC change current, TIA gain, number of pulses slightly and run the test | PPG test without AGC is not affected by AGC; One can change the value of all electrical parameters of coolidge | PPG tests executes; all the changes made to electrical parameters are stored correctly | Needs to be implemented |  | Praptee and Nitin | M | Functional/System |  |  | Check with Aparna for CLI command
Temperature Tests
25 | Temperature Accuracy | All | All |  |  |  |  | check 'Temperature' sheet of this excel for details | Temperature value | +/-' 0.1 degree Celcius | Done | NA | NA | M | Performance
Bluetooth Tests
26 | BLE Functionality | All | Few |  |  |  |  | Collect data for the four test cases from 'Use Cases' tab | BLE works without packet loss at the corner cases drived from use cases | No Packet loss recorded | Done |  | Nitin | H | Performance |  |  | Check with Aparna
27 | BLE Reliability at different distances | Few | Few |  |  |  |  | Perform Ping test at incremental distances | BLE reliability at longer distances | BLE stays connected and no packet loss observed | Done |  | Nitin | M | Performance |  | 2020-05-19
28 | BLE Reliability | Few | Few |  |  |  |  | Connect, disconnect and reconnect watch with ble dongle repeat 10 times | BLE should connect everytime | No connection issue | Not a automated test yet |  | Nitin | L | Performance
Power Tests
29 | Battery test: Current consmptoin in streaming mode | NA | All |  |  |  |  | check 'Battery' sheet of this excel for details | Current consumption is within limit |  | Done |  | Praptee, Aparna and Nitin | H | Performance
30 | Battery test: Current consmptoin in BLE advertising mode | [8 empty] | <= 17mA
31 | Battery test: Current consmptoin in hibernate mode | [14 empty] | Mode not Available
32 | Battery test: Current consmptoin in charging mode | [8 empty] | <=99mA
Bio-Z Tests
33 | AD5940 switch (DG2502) | All | All |  |  |  |  | Check 'BioZ' tab for details | Is AD5940 switch working | Output should match the input when swtich on and show noise when switch off amplitude < -80dB | Need to add auto data analysis in test script |  | Nitin | H | Functional | Waitin for the logging issue to be resolved; Test framwork is complete just need to validate it
34 | AD5940 accuracy | All | All |  |  |  |  | Enable 5940 Switch and obtain value of a Known Impedance; contact impedance zero Ohm. Test with 50 Ohm, 100 Ohm and 1kOhm | Check the accuracy of Impedance measurement | Magnitude accuracy within 3% | Test yet to be implemented. Waiting on getting the firmware app ready |  | Praptee & Nitin | H | Performance
35 | AD5940 repeatabiliy | Few | All |  |  |  |  | Enable 5940 Switch and obtain value of a Known Impedance; contact impedance zero Ohm. Test with 50 Ohm, 100 Ohm and 1kOhm. Repeat 10 times | Is AD5940 output repetable | Magnitude standard deviation < 0.1Ohm | Test yet to be implemented. Waiting on getting the firmware app ready |  | Praptee & Nitin | H | Functional
36 | AD5940 accuracy with non-ideal contact impedance | Few | All |  |  |  |  | Enable 5940 Switch and obtain value of a Known Impedance; contact impedance 1kOhm, 2kOhm and 3kOhm | AD5940 should be able to produce accurate enough result with imbalance in contact impedance | Magnitude accuracy within 3% | Test yet to be implemented. Waiting on getting the firmware app ready |  | Praptee & Nitin | H | Performance
Functional Tests
37 | Memory Read/Write Test | All | All |  |  |  |  | Enable the memory and write random words to random location; Read back these words | Check if the memory is able to store the values correctly | Words should match | Done |  | Nitin | M | System/Firmware |  | 2020-05-20
38 | SPI communication test | All | All |  |  |  |  | Read Device ID of each device AD5940, ADPD4000, ADP5360, ADXL362, AD7156 | Check if SPI linese are connected correctly and protocal is working | ID should match the device | Not yet implemented |  | Nitin | L | System/Firmware
39 | Connectivity and communication | All | All |  |  |  |  | connect the dut to computer over cable and get connected successfully message | Check if the device connects with the computer over usb | Dut should connect to every computer/laptop | Kind of implemented needs to implement a check |  | Nitin | M | System |  | 2020-05-21
40 | Button | All | few |  |  |  |  | press the buttons and look for appropriate response | Button are responding to the firmware commands | correct response is recorded everytime | Not yet implemented |  | Praptee, Aparna and Nitin | M | System |  | 2020-05-22
Cap Touch Tests
41 | AD7156 | All | All |  |  |  |  | Use a metal plate to generate an output signal. Check the output with and without a metal plate placed below the watch | Capacitive sensing is working properly | Output is observed with the plate and low level noise without the plate | Not yet implemented |  |  | L | Functional
Additional Firmware and miscellaneous tests
42 | Firmware version | All | All |  |  |  |  | Read and store the firmware version information | For information and debugging | NA | Done |  | Nitin | M | System |  | 2020-05-19
43 | ADPD register read test | None | All |  |  |  |  | Check-read-modify-read and check for the correctness of register value | To make sure ADPD register read does not have any issue | Check if the registers can be modified | Not yet implemented |  |  | L | Firmware
44 | HR test | All | All |  |  |  |  | Test the HR value for PPG signal using the HR algorithm | To check the signal chain continuity from Hardware to algorithm and back to the system | Check if the HR value are within the normal range | Not yet implemented |  |  | L | Firmware/software
45 | Display test | All | Few |  |  |  |  | Issue command to display all white,R,G,B and black, to check whether have dead pixel. | Check the integrity of the display | There should not be any dead pixel | Not yet implemented |  | Nitin | M | System/functional |  | 2020-05-22

## Performance
Functions | Test | Watch PCB controls | DAC0 | DAC1 | GPIOs | Procedure | Pass criteria | Pass/Fail | Comments
ECG | AD8233 Switch functionality | Turn off AD5940 and ADPD4002 Switches; Turn on and off the 8233 Switch and check if AD8233 is disconneted from the electrodes | Sinwave(pos) freq:5Hz, amp= 0.1Vpp | Sinwave(neg) freq:5Hz, amp= 0.1Vpp | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Close DG2502 switch and collect data for 10 sec Open DG2502 switch and collect data for 10 sec | Close switch :amplitude matches expected ampllitude (+-10%) Open switch: ECG output is lost (no power at 5Hz). Inputs float so there may be energy at 50Hz |  | This test should be performed before AD5940 Switch Test
| ECG Signal quality with ECG simulator | Turn off AD5940 and ADPD4002 Switches; Turn on the 8233 Switch and collect AD5940 output data | - | - | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:1-2 JP8:2-3 | Display the waveform and record the data for further examination | Take FFT should not see any 50,60Hz noise
| ECG Signal quality on body | Turn off AD5940 and ADPD4002 Switches; Turn on the 8233 Switch and collect AD5940 output data | - | - | SNOISE1:0 SNOISE2: 0 | Connect one end wet electrodes to Pins 8,6 and 5 of JP14 and place the wet electrodes on left/right wrist and two fingers of the other hand. Display the waveform and save the data for further examination | Match the data with the best waveform recorded by Gen 3 watch
Accelerometer | Self Test for ADXL362 | Run the self test for ADXL362 | - | - | - | Run the ADXL362 self test | Pass criteria as per datasheet
PPG | Dark Test | Turn on ADPD to measure PPG | - | - | - | Place a refective surface over the LEDs and measure the output at 1 pulse 0mA current | dBm >= 85
| Test without AGC | Turn on ADPD to measure PPG | - | - | - | Place the finger/wrist on the optical window and display PPG using wavetool | Good signal (close to finger clip) On Praptee, Shane, Supriya, Dan, Foroohar
| Test with AGC | Turn on ADPD to measure PPG | - | - | - | Place the finger/wrist on the optical window and display PPG using wavetool | Good signal (close to finger clip) On Praptee, Shane, Supriya, Dan, Foroohar
Temperature | Accuracy | Place a metal bar with (palladin) over the RTD and change the temperature | - | - | -
BLE | Functionality | Connect BLE to Watch PCB and collect data |  |  |  | Collect data for an hour of ECG, PPG, Temperature and ADXL at all the frequencies to be demod at CES | BLE stays connected and no packet lose reported
| Reliability | Collect data over BLE from N channels (as recommended by Firmware team) over night | - | - | - | Collect PPG, Temperature and ADXL data for an hour and ECG data for another hr | No data lost
Battery Testing |  | Follow Fenda's tests

## ECG
Test | Watch PCB Controls | DAC0 | DAC1 | GPIO | Procedure | Priority | Pass Criteria | Pass/Fail | Comments
DG2502 | Turn off AD5940 and ADPD4002 Switches; Turn on and off the 8233 Switch and check if AD8233 is diconneted from the electrodes | Sinwave(pos) freq:5Hz, amp= 0.05Vpp | Sinwave(neg) freq:5Hz, amp= 0.05Vpp | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Close DG2502 switch and collect data for 10 sec Open DG2502 switch and collect data for 10 sec | H | Close switch :amplitude matches expected ampllitude (+-10%) Open switch: ECG output is lost (no power at 5Hz). Inputs float so there may be energy at 50Hz |  | This test should be performed before AD5940 Switch Test amplitude < -80dB
AD8233 Noise | Enale 8233 switch, connect the contacts to ECG out of interface board, inject know signal value. Collect ADC output data for 61 seconds | NA | NA | SNOISE1:0 SNOISE2: 1 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Connect all the electrodes and RLD will provide 1.1Vdc; collect 50 second signal data and Calculate Vpp for the 11 second signal | H | Noise < 50uVpp |  | IEC2-27: Input Noise < 30uVpp for at least 10s measurement. Can do shorter test for production? Dan: Yes; not sure how long
AD8233 Bandwidth | Enable 8233 Switch and connect the contacts to the ECG out of interface board and check the Bandwidth Attenuation from 0.5Hz to 40Hz | Sinwave(pos) freq:40Hz - 0.5Hz, amp= 0.15Vpp | Sinwave(neg) freq:40Hz - 0.5Hz, amp= 0.15Vpp | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Test Frequencies- 40Hz, 30Hz, 20Hz, 10Hz, 5Hz, 0.5Hz: ECG Common: 1.1Vpp, DC; Vpp at each frequency should match the theortical value | H | Monitor BW 0.34Hz - 38Hz Difference <TBD |  | IEC2-27: Frequency and impulse response BW: 0.67Hz to 40Hz with 71% to 110% of output obtained at 5Hz
AD8233 Bandwidth
AD8233 dynamic range | Enable the 8233 Switch and connect the contacts to the ECG out of interface board and check if the AD8233 gain is "constant" | Sinwave(pos) freq:5Hz, amp= 0.1Vpp-0.5Vpp | Sinwave(neg) freq:5Hz, amp= 0.1Vpp-0.5Vpp | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Test amplitudes: 0.5Vpp, 0.1Vpp & 0.2Vpp & 0.5Vpp Measure Vrms at 5Hz for each amplitude | H | gain tolerance(150) +-10% |  | IEC2-27: Input dynamic range +-10% over specified DC offset
AD8233 Repeatability | Enable the 8233 Switch and connect the contacts to the ECG out of interface board and check if the AD8233 repetability is good | Sinwave(pos) freq:5Hz, amp= 0.1Vpp | Sinwave(neg) freq:5Hz, amp= 0.1Vpp | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Test at 5Hz, 1mVpp . Repeat 10 times. Calculate Std | H | SD < TBD |  | IEC2-27: Accuracy of signal reproduction +-20% or +-100uV
AD8233 CMRR | Enable the 8233 Switch and connect the contacts to ECG out of the interface board and measure how good is the AD8233 eliminating the 50-60 Hz noise | Sinwave(pos) freq:50Hz, amp= 0.1Vpp | Sinwave(pos) freq:50Hz, amp= 0.1Vpp | testa: SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 testb: SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 0 ECG_POSSEL: 1 testc: SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 0 JP7:2-3 JP8:1-2 | Check 20LOG( (amplitude_ecg_codes/gain)/ amplitude_input signal ). Test with no electrode imbalance . Target >80dB - Repeat with 51k¦¦47nF imbalance in one electrode. Target > 40dB | H | CMRR>80dB (no electrode imbalance) CMRR>40dB 51k¦¦47nF electrode imbalance) |  | IEC2-27: CMRR
AD8233 input impedance | Enable the 8233 Switch and connec the contacts to the ECG out of interface board | Sinwave(pos) freq:5Hz, amp= 0.1Vpp | Sinwave(neg) freq:5Hz, amp= 0.1Vpp | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Replace R25 and R26 by 1M resistor. Check output amplitude. Add 1M in series with each input and check change in output amplitude. Calculate Zin | H | >10M |  | IEC2-27: Input Impedance >2.5MoHm

## ADPD
Test | Watch PCB Controls | DAC0 | DAC1 | GPIO | Procedure | Priority | Pass Criteria | Pass/Fail | Comments
DG2502 for ADPD4002 | Test Switch for ADPD4002 | Sinwave(pos) freq:5Hz, amp= 0.1Vpp | Sinwave(neg) freq:5Hz, amp= 0.1Vpp | SNOISE1:0 SNOISE2: 0 ECG_NEGSEL: 0 ECG_POSSEL: 0 JP7:2-3 JP8:1-2 | Close DG2502 switch and collect data for 10 sec Open DG2502 switch and collect data for 10 sec | H | Close switch :amplitude matches expected ampllitude (+-10%) Open switch: ECG output is lost (no power at 5Hz). Inputs float so there may be energy at 50Hz
Read device ID | Read the ADPD ID | NA | NA | SNOISE1:0 SNOISE2: 0 ECG_NEGSEL: 0 ECG_POSSEL: 0 JP7:2-3 JP8:1-2 | Read ADPD4002 ID | H | Correct ID is read back

## PPG
Test | Watch PCB Controls | DAC0 | DAC1 | GPIO | Procedure | Priority | Pass Criteria | Pass/Fail | Comments
Dark current | Set all LED current values to 0mA; frequency to 100Hz and No. of pulses to 1; turn on both the channels | NA | NA | SNOISE1:0 SNOISE2: 0 ECG_NEGSEL: 0 ECG_POSSEL: 0 JP7:2-3 JP8:1-2 | Collect the data for 10sec and check the standard deviation for both the channels | H | std < 2
LED Drivers | Place the finger clip board on the HR board at some distance and turn LEDs one at a time set the current to 2mA, number of pulses 16 | NA | NA | SNOISE1:0 SNOISE2: 0 ECG_NEGSEL: 0 ECG_POSSEL: 0 JP7:2-3 JP8:1-2 | Collect the data for 10sec on the finger clip board PD and match the current value | H | Match the value of the PD output to the input current
PD | Place the finger clip board on the HR board at some distance and turn on green LED on the finger clip board to 2mA, 16 pulses at a time set the current to 2mA, number of pulses 16 and turn on one channel at a time on the HR board | NA | NA | SNOISE1:0 SNOISE2: 0 ECG_NEGSEL: 0 ECG_POSSEL: 0 JP7:2-3 JP8:1-2 | Collect the data for 10sec on HR board and match the input current to PD output | H | Match the value of the PD output to the input current

## BioZ
Test | Watch PCB Controls | DAC0 | DAC1 | GPIO | Procedure | Priority | Pass Criteria | Pass/Fail | Comments
Test Switch | Turn off AD8233 and ADPD4002 Switches; Turn on and off the 5940 Switch and check if AD5940 is diconneted from the electrodes | Sinwave(pos) freq:5Hz, amp= 0.1Vpp | Sinwave(neg) freq:5Hz, amp= 0.1Vpp | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Close DG2502 switch and collect data for 10 sec Open DG2502 switch and collect data for 10 sec | H | Output should match the input when swtich on and show noise when switch off amplitude < -80dB
AD5940 accuracy | Enable 5940 Switch and Implement frequncy Sweep and obtain value of a Known Impedance; contact impedance zero Ohm | Sinwave(pos) amp= 0.1Vpp; Sweep frequency | Sinwave(neg) amp= 0.1Vpp; Sweep frequency | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Sweep across 5KHz, 25KHz, 50KHz, 100KHz, 150KHz, 180KHz, 190KHz and 200KHz. Test with 50 Ohm, 100 Ohm and 1kOhm | H | Magnitude accuracy within 3% Phase accuracy within 6 degrees
AD5940 repeatabiliy | Enable 5940 Switch and Implement frequncy Sweep and obtain value of a Known Impedance; contact impedance 0 Ohm | Sinwave(pos) amp= 0.1Vpp; Sweep frequency | Sinwave(neg) amp= 0.1Vpp; Sweep frequency | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Sweep across 5KHz, 25KHz, 50KHz, 100KHz, 150KHz, 180KHz, 190KHz and 200KHz. Test with 50 Ohm, 100 Ohm and 1kOhm; repeat each reading 10 times | H | Magnitude standard deviation < 0.1Ohm and Phase standard variation <0.5 degree
AD5940 repeatabiliy
AD5940 accuracy with non-ideal contact impedance | Enable 5940 Switch and Implement frequncy Sweep and obtain value of a Known Impedance; contact impedance 1kOhm, 2kOhm and 3kOhm | Sinwave(pos) amp= 0.1Vpp; Sweep frequency | Sinwave(neg) amp= 0.1Vpp; Sweep frequency | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Sweep across 5KHz, 25KHz, 50KHz, 100KHz, 150KHz, 180KHz, 190KHz and 200KHz. Test with 100 Ohm, 50 Ohm and 1kOhm | H | Magnitude accuracy within 3% Phase accuracy within 6 degrees |  | Pass criteria might change
AD5940
AD5940 accuracy with contact impedance mismatch | Enable 5940 Switch and Implement frequncy Sweep and obtain value of a Known Impedance; contact impedance set F+/S+/F-/S- to 2/2/2/1 kOhm | Sinwave(pos) amp= 0.1Vpp; Sweep frequency | Sinwave(neg) amp= 0.1Vpp; Sweep frequency | SNOISE1:1 SNOISE2: 0 ECG_NEGSEL: 1 ECG_POSSEL: 1 JP7:2-3 JP8:1-2 | Sweep across 5KHz, 25KHz, 50KHz, 100KHz, 150KHz, 180KHz, 190KHz and 200KHz. Test with 100 Ohm, 50 Ohm and 1kOhm | H | Magnitude accuracy within 3% Phase accuracy within 6 degrees (pass criteria might change) |  | Change the mismatch impedance to all the four contacts and repeat the test
AD5940

## Temperature
Test | Watch PCB Controls | DAC0 | DAC1 | GPIO | Procedure | Priority | Pass Criteria | Pass/Fail | Comments
Accuracy /Reliability | Place the Palladin above the temp sensor increase and decrease the temperature in steps and match the value from ADPD output | NA | NA | NA |  | H | Temperature value should match within 1 degree

## Battery
Test | Watch PCB Controls | DAC0 | DAC1 | GPIO | Procedure | Priority | Pass Criteria | Pass/Fail | Comments
Charging Current Consumption | enter PCB into specified mode. | NA | NA | NA | Connect Keithley (V/ Source/Meter) to PCB |  | Verify that Charging Current Limits [-45mA:-14mA]
Static Current Consumption | enter PCB into specified mode. | NA | NA | NA | Connect Keithley (V/ Source/Meter) to PCB |  | Verify that Static Current Limits [1mA:5mA]
Hibernate Current Consumption | enter PCB into specified mode. | NA | NA | NA | Connect Keithley (V/ Source/Meter) to PCB |  | Verify that Hibernate Current Limits [1mA:5mA]
Sweep Stage Current Consumption | enter PCB into specified mode. | NA | NA | NA | Connect Keithley (V/ Source/Meter) to PCB |  | Verify Current Limits |  | Frequency sweep for AD5940
Streaming Stage Current Consumption | enter PCB into specified mode. | NA | NA | NA | Connect Keithley (V/ Source/Meter) to PCB |  | Verify Current Limits |  | PPG @50Hz, ADXL @50Hz, EDA @ 100Hz and Skin temp @1Hz with Display ON and OFF
BLE advertising Stage Current Consumption | enter PCB into specified mode. | NA | NA | NA | Connect Keithley (V/ Source/Meter) to PCB |  | Verify Current Limits

## Schedule
Sr.No# |  |  | comments | Resourses
| Performance Tests | 2020-01-23 |  | Nitin, Praptee and Shane; Dan as needed
1 | AD8233 Switch functionality
2 | ECG Signal quality with ECG simulator
3 | ECG Signal quality on body
4 | Self Test for ADXL362 |  | Will be added by Wednesday
5 | PPG Dark Test
6 | PPG Test without AGC
7 | PPG Test with AGC |  | depends on Firmware
8 | Temperature Accuracy |  | will be done as part of temperature testing
9 | BLE Functionality | 2020-01-31 | Need to discuss test parameters with the team
10 | BLE Reliability
11 | Battery
| ECG for AD8233 | 2020-01-31
12 | 8233 switch (DG2502)
13 | AD8233 input Noise |  | Will be ready by Monday
14 | AD8233 Bandwidth
15 | AD8233 dynamic range
16 | AD8233 Repeatability
17 | AD8233 CMRR
18 | AD8233 input impedance
| BioZ for AD5940 | 2-3days | depends on Firmware to release the impedance test app
19 | 5940 switch (DG2502)
20 | AD5940 accuracy
21 | AD5940 repeatabiliy
22 | AD5940 accuracy with non-ideal contact impedance
23 | AD5940 accuracy with contact impedance mismatch
| Battery | 2020-02-05 | Will be ready by Wednesday
24 | Charging Current Consumption
25 | Static Current Consumption
26 | Hibernate Current Consumption
27 | Sweep Stage Current Consumption
28 | Streaming Stage Current Consumption
29 | BLE advertising Stage Current Consumption
| Temperature | 2020-01-31
30 | Accuracy /Reliability
| PPG | 2-3days | do we need these tests ??
31 | Dark current
32 | LED Drivers
33 | PD
| ECG for ADPD4000 | 2-3days | depends on Firmware to release the ADPD BioZ test app
34 | 8233 switch (DG2502)
35 | AD8233 input Noise
36 | AD8233 Bandwidth
37 | AD8233 dynamic range
38 | AD8233 Repeatability
39 | AD8233 CMRR
40 | AD8233 input impedance
| BioZ for ADPD4000 | 2-3days | depends on Firmware to release the ADPD ECG test app
41 | 5940 switch (DG2502)
42 | AD5940 accuracy
43 | AD5940 repeatabiliy
44 | AD5940 accuracy with non-ideal contact impedance
45 | AD5940 accuracy with contact impedance mismatch

## Use cases
Case1: PPG Specific
Primitives | Sensor Operating Condition
PPG @500 Hz | 32 Bit Mode
Accelerometer @50 Hz | 8g Mode
Skin Temp @1 Hz | (see measurement details)
Display | On or Off
Battery Life (streaming over BLE) | TBD (Display = Off)
Battery Life (recording to Flash) | TBD (Display = Off)
Logging Default | ON
Low Touch Wrist | Turns logging OFF after 5 seconds off the wrist
Turns logging ON after 5 seconds on the wrist
Low Touch Watch Face | Turns on Backlight (Tap)…. Remains on for 15 seconds
Case 2: PPG + EDA
Primitives | Sensor Operating Condition
PPG @100 Hz | 32 Bit Mode
Accelerometer @50 Hz | 8g Mode, match XL ODR to PPG
EDA at @30 Hz | (see measurement details)
Skin Temp @1 Hz | (see measurement details)
Display | On or Off
Battery Life (streaming over BLE) | TBD (Display = Off)
Battery Life (recording to Flash) | TBD (Display = Off)
Low Touch Wrist | Turns logging OFF after 5 seconds off the wrist
Turns logging ON after 5 seconds on the wrist
Low Touch Watch Face | Turns on Backlight (Tap)…. Remains on for 15 seconds
Case 3: PPG + ECG
Primitives | Sensor Operating Condition
PPG @100 Hz | 32 Bit Mode
Accelerometer @50 Hz | 8g Mode, match XL ODR to PPG
ECG @250Hz | (see measurement details)
Skin Temp @1 Hz | (see measurement details)
Battery Life (streaming over BLE) | TBD (Display = Off)
Battery Life (recording to Flash) | TBD (Display = Off)
Low Touch Wrist | Turns logging OFF after 5 seconds off the wrist
Low Touch Watch Face | Turns on Backlight (Tap)…Remains on for 15 seconds
Turns on ECG application after 2 seconds of being held
Turns off ECG application after 2 seconds of being released
Case 4: ECG
Primitives | Sensor Operating Condition
PPG @50 Hz | 32 Bit Mode
Accelerometer @50 Hz | 8g Mode, match XL ODR to PPG
ECG @1000Hz | (see measurement details)
Skin Temp @1 Hz | (see measurement details)
Battery Life (streaming over BLE) | TBD (Display = Off)
Battery Life (recording to Flash) | TBD (Display = Off)
Low Touch Wrist | Turns logging OFF after 5 seconds off the wrist
Low Touch Watch Face | Turns on Backlight (Tap)…Remains on for 15 seconds
Turns on ECG application after 2 seconds of being held
Turns off ECG application after 2 seconds of being released

## OLD_Test_list
Test# | Application | Block | Paramter | Test Type | Test Type 2 | Procedure | Test Setup | Methodology | Priority | Pass Criteria | Status | Ownership | Comments | Test Level
1 | System | Uart<->USB | Communication | Functional | Self Test | Send and receive data via UART | Diagram 1 | Send and receive a specific word by Uart | L | returns same word |  |  |  | Final Functional
2 | System | nRf52840 | HF XTAL, PLL and LF OSC | Functional | Self Test | Set the internal Clock and toggle a GPIO in order to collect data on oscilloscope | Diagram 1 | Change to other available frequency and repeat the test | L | Clocks are stable |  |  | Get realibility report from supplier. Can crosscheck between crystal oscillators | Charaterization
4 | System | Memory | Enable/Disable | Functional | Self Test | Check the access to memory when enabled and disabled | Diagram 1 | Check the status register during power off for low volatge as per datasheet | L | Words should not match |  |  | Uses QPSI. Check voltage via register read during power off. | Firmware test
5 | System | Memory | Read/Write | Functional | Self Test | Check if the memory is able to store the values correctly | Diagram 1 | Enable the memory and write random words to random location; Read back these words | M | Words should match |  |  | Checker board test | Firmware test
6 | System | AD5940 | SPI communication | Functional | Self Test | Read ID | Diagram 1 | Read the AD5940 ID | M | Correct ID is read back |  |  | Already checked from drivers | Final Functional
7 | System | AD5940 | GPIOs | Functional | Self Test | Test the AD5940 GPIOs for register read/write | Diagram 1 | Toggle the GPIOs High and low and check the levels | L | Levels matches the setting |  |  |  | Final Functional
8 | BioZ | DG2502 for AD5940 | Test Switch | Functional | Signal chain integrity test | Turn off AD8233 and ADPD4002 Switches; Turn on and off the 5940 Switch and check if AD5940 is diconneted from the electrodes | Diagram 1 | Use AIN1 and AIN2 to inject a 5 Hz sinusoidal waveform; amp: 0.3Vpp; offset: 1.8V and common mode DC 1.1V sigal via Interface board in AD5940 and check the output | H | Output should match the input when swtich on and show noise when switch off amplitude < -80dB |  |  |  | Final Functional
9 | BioZ | AD5940 and DG2502 | AD5940 accuracy | Performance |  | Enable 5940 Switch and Implement frequncy Sweep and obtain value of a Known Impedance; contact impedance zero Ohm | Diagram 1 | Sweep across 5KHz, 25KHz, 50KHz, 100KHz, 150KHz, 180KHz, 190KHz and 200KHz. Test with 100 Ohm, 510 Ohm and 1kOhm | H | TBD |  |  | For production test use only few points | Characterization
10 | BIOZ | AD5940 and DG2502 | AD5940 accuracy with non-ideal contact impedance | Performance |  | Enable 5940 Switch and Implement sweep in frequency and obtain the value of a known impedance; contact impedance > 0 Ohm | Diagram 1 | Sweep across 5KHz, 25KHz, 50KHz, 100KHz, 150KHz, 180KHz, 190KHz and 200KHz. Test with 100 Ohm, 510 Ohm and 1kOhm | H | TBD |  |  | Not needed for production | Characterization
| BIOZ |  | AD5940 | Performance
11 | BIOZ | AD5940 and DG2502 | AD5940 repeatabiliy | Performance |  | Enable 5940 Switch and Implement sweep in frequency and obtain the value of a known impedance | Diagram 1 | Post process the data from previous measurement and analyze the standard deviation | H | TBD |  |  | Repeat each meausrement 20 or so times for characterization. Repeat only a few times for produciton. | Characterization
| BIOZ |  | AD5940 repeatabiliy | Performance
12 | System | ADXL362 | SPI communication | Functional | Self Test | Read ID | Diagram 1 | Read the ADXL362 ID | L | Correct ID is read back |  |  |  | Final Function
13 | System | ADXL362 | Self Test | Functional | Self Test | Run the ADXL362 self test | Diagram 1 | Run ADXL362 self test | H | Pass criteria as per datasheet |  |  |  | FF/Module
DIAGRAM 1
14 | DG2502 for AD8233 | Test Switch For AD8233 | DG2502 | Functional |  | Turn off AD5940 and ADPD4002 Switches; Turn on and off the 8233 Switch and check if AD8233 is diconneted from the electrodes | Diagram 1 | Close switch. Inject sine wave 5Hz, 1mVpp, Vdc = 1.1V, and check amplitude matches expected ampllitude (+-10%) Open switch. ECG output is lost (no power at 5Hz). Inputs float so there may be energy at 50Hz | H | amplitude < -80dB |  |  | This test should be performed before AD5940 Switch Test | Final Functional
15 | ECG | AD8233 and DG2502 | AD8233 Noise | Performance |  | Enale 8233 switch, connect the contacts to ECG out of interface board, inject know signal value. Collect ADC output data for 61 seconds | Diagram 1 | ECG Differential : Offset - 1.8V, Amplitude- 0.01Vpp, Frequency- 0.1Hz ECG Common: 1.1Vpp, DC; Clip 50 second signal data and Calculate Vpp for the 11 second signal | H | Noise < 50uVpp |  |  | IEC2-27: Input Noise < 30uVpp for at least 10s measurement. Can do shorter test for production? Dan: Yes; not sure how long | Characterization |  | Diagram 1
Performance
16 | ECG | AD8233 and DG2502 | AD8233 Bandwidth | Performance |  | Enable 8233 Switch and connect the contacts to the ECG out of interface board and check the Bandwidth Attenuation from 0.5Hz to 100Hz | Diagram 1 | ECG Differential : Offset - 1.8V, Amplitude- 0.3Vpp, Frequency- 100Hz, 70Hz, 40Hz, 30Hz, 20Hz, 10Hz, 5Hz, 0.5Hz: ECG Common: 1.1Vpp, DC; Vpp at each frequency should match the theortical value | H | Monitor BW 0.34Hz - 38Hz Difference <TBD |  |  | IEC2-27: Frequency and impulse response BW: 0.67Hz to 40Hz with 71% to 110% of output obtained at 5Hz | Characterization
| ECG | AD5940 & AD8233 & DG2502 | AD8233 Bandwidth | Performance
17 | ECG | AD8233 and DG2502 | AD8233 dynamic range | Performance |  | Enable the 8233 Switch and connect the contacts to the ECG out of interface board and check if the AD8233 gain is "constant" | Diagram 1 | Test at 5Hz, 0.5mVpp, 1mVpp & 2mVpp & 5mVpp (5mVpkpk*296=1.48V=+-0.74V) Measure Vrms at 5Hz for each amplitude | H | gain tolerance +-10% |  |  | IEC2-27: Input dynamic range +-10% over specified DC offset | Characterization
18 | ECG | AD8233 and DG2502 | AD8233 Repeatability | Performance |  | Enable the 8233 Switch and connect the contacts to the ECG out of interface board and check if the AD8233 repetability is good | Diagram 1 | Test at 5Hz, 1mVpp . Repeat 10 times. Calculate Std Dev | H | SD < TBD |  |  | IEC2-27: Accuracy of signal reproduction +-20% or +-100uV
19 | ECG | AD8233 and DG2502 | AD8233 CMRR | Performance |  | Enable the 8233 Switch and connect all the contacts to either P or N ECG out of the interface board and measure how good is the AD8233 eliminating the 50-60 Hz noise | Diagram 1 | Drive a 50Hz sinewave Vdc=1.1V Vac=0.3Vpp. Check 20LOG( (amplitude_ecg_codes/gain)/ amplitude_input signal ). Test with no electrode imbalance . Target >80dB - Repeat with 51k¦¦47nF imbalance in one electrode. Target > 40dB | H | CMRR>80dB (no electrode imbalance) CMRR>40dB 51k¦¦47nF electrode imbalance) |  |  | IEC2-27: CMRR
20 | ECG | AD8233 and DG2502 | AD8233 input impedance | Performance |  | Enable the 8233 Switch and connec the contacts to the ECG out of interface board | Diagram 1 | Test at 5Hz, 1mVpp. Check output amplitude. Add 1M in series with each input and check change in output amplitude. Calculate Zin | H | >10M |  |  | IEC2-27: Input Impedance >2.5MoHm
21 | ECG | AD8233 | AD8233 RLD | Performance |  | Check RLD bandwidth and gain with the scope | Diagram 1 | Drive sinewave 10mVpp at different freqeuncies and check volatge at RLD. Comapre to theoretical values. Frequency and Amplitudes to be confirmed. Internal R in AD8233 specifed at +-30% | H | TBD |  |  | Characterize at a couple points | FF/Module
22 | System | ADPD4002 | SPI communication | Functional | Self Test | Read ID | Diagram 1 | Read ADPD4002 ID | L | Correct ID is read back |  |  |  | Final Functional
23 | System | ADPD4002 | GPIOs | Functional | Self Test | Test the ADPD4002 GPIOs are soldered correctly | Diagram 1 | Toggle the GPIOs High and low and check the levels | L | Levels matches the setting |  |  |  | Final Functional
24 | System | ADPD4002 | ADC Vref | Functional | Self Test | Use Vc2 connected at ET1/EB2 and measure the value | Diagram 1 | TBD | M | as per datasheet |  |  |  | Characterization
25 | System | ADPD4002 and LEDs | Chcek the led Green and Red LED on/off | Functional | User Interface | Toggle the LEDs on and off | Diagram 1 | Toggle LEDs on/off and make a visual observation if they are as per setting | H | LEDs turn On and Off |  |  |  | Final Functional
26 | System | ADPD4002 and LEDs | Check the IR LED on/off | Fucntional | User Interface | Use a black box to cover the LEDS and turn on the IR led collect the reflection on a PD connected to the black box | Diagram 1 | Pass the known excitation to the IR LED and oberve the equivalent output at the PD | H | PD output should match calculated value |  |  |  | Final Functional
27 | System | ADPD4002 and PD | Check the dark current and light on the PD | Functional | User Interface | Use a black box to cover the LEDS and turn on/off the green LED connected to the box; collect the dark current and light reflection on a PD | Diagram 1 | Pass the known exciation current through the green LED on the box and measure the dark current and light reflection | H | PD output should match the datasheet |  |  |  | Final Functional
28 | DG2502 for ADPD4002 | Test Switch for ADPD4002 | DG2502 | Functional |  | Turn off AD8233 and AD5940 Switches; Turn on and off the 4002 Switch and check if ADPD4002 is diconneted from the electrodes | Diagram 1 | Close switch. Inject sine wave 5Hz, 1mVpp, Vdc = 1.1V, and check amplitude matches expected ampllitude (+-10%) Open switch. ECG output is lost (no power at 5Hz). Inputs float so there may be energy at 50Hz | M | amplitude < -80dB |  |  |  | Final Functional
29 | BioZ | ADPD4002 and DG2502 | Accuracy | Performance |  | Enable 4002 Switch and Implement frequncy Sweep and obtain value of a Known Impedance; contact impedance zero Ohm | Diagram 1 | Sweep across 5KHz, 25KHz, 50KHz, 100KHz, 150KHz, 180KHz, 190KHz and 200KHz. Test with 100 Ohm, 510 Ohm and 1kOhm | M | TBD |  |  |  | Characterization/FF- only few points
30 | BIOZ | ADPD4002 and DG2502 | ADPD4002 accuracy with non-ideal contact impedance | Performance |  | Enable 5940 Switch and Implement sweep in frequency and obtain the value of a known impedance; contact impedance > 0 Ohm | Diagram 1 | Sweep across 5KHz, 25KHz, 50KHz, 100KHz, 150KHz, 180KHz, 190KHz and 200KHz. Test with 100 Ohm, 510 Ohm and 1kOhm | M | TBD |  |  |  | Characterization/FF- only few points
| BIOZ |  | AD5940 | Performance
31 | BIOZ | ADPD4002 and DG2502 | ADPD4002 repeatabiliy | Performance |  | Enable 4002 Switch and Implement sweep in frequency and obtain the value of a known impedance | Diagram 1 | Post process the data from previous measurement and analyze the standard deviation | M | TBD |  |  |  | Characterization/FF- only few points
| BIOZ |  | AD5940 repeatabiliy | Performance
32 | ECG | ADPD4002 and DG2502 | ADPD4002 Noise | Performance |  | Enale 4002 switch and using interface board produce know value of differential and common mode signal. Collect ADC output data for 61 seconds | Diagram 1 | ECG Differential : Offset - 1.8V, Amplitude- 0.01Vpp, Frequency- 0.1Hz ECG Common: 1.1Vpp, DC; Clip 50 second signal data and Calculate Vpp for the 11 second signal | M | Noise < 50uVpp |  |  |  | Characterization
Performance
33 | ECG | ADPD4002 and DG2502 | ADPD4002 Bandwidth | Performance |  | Enable 4002 Switch and connect the contacts to the ECG out of interface board and check the Bandwidth Attenuation from (same)?? | Diagram 1 | ECG Differential : Offset - 1.8V, Amplitude- 0.3Vpp, Frequency- 40Hz, 30Hz, 20Hz, 10Hz, 5Hz, 0.5Hz, 0.1Hz : ECG Common: 1.1Vpp, DC; Vpp at each frequency should match the theortical value | M | Monitor BW 0.34Hz - 38Hz Difference <TBD
| ECG | AD5940 & AD8233 & DG2502 | AD8233 Bandwidth | Performance
34 | ECG | ADPD4002 and DG2502 | ADPD4002 Dynamic range | Performance |  | Enable the 4002 Switch and connect the contacts to the ECG out of interface board and check if the ADPD4002 gain is "constant" | Diagram 1 | Test at 5Hz, 0.5mVpp, 1mVpp & 2mVpp & 5mVpp (5mVpkpk*296=1.48V=+-0.74V) Measure Vrms at 5Hz for each amplitude | M | gain tolerance ??
|  | AD5940 & AD8233 & DG2502 | AD8233 Bandwidth | Performance
35 | ECG | ADPD4002 and DG2502 | ADPD4002 Repeatability | Performance |  | Enable the 4002 Switch and connect the contacts to the ECG out of interface board and check if the ADPD4002 repetability is good | Diagram 1 | Test at 5Hz, 1mVpp . Repeat 10 times. Calculate Std Dev | M | SD < TBD
| ECG | AD8233 and DG2502 | AD8233 Repeatability | Performance |  | Enable the 8233 Switch and connect the contacts to the ECG out of interface board and check if the AD8233 repetability is good |  | Test at 5Hz, 1mVpp . Repeat 10 times. Calculate Std Dev |  | SD < TBD
36 | ECG | ADPD4002 and DG2502 | ADPD4002CMRR | Performance |  | Enable the4002 Switch and connect all the contacts to either P or N ECG out of the interface board and measure how good is the ADPD4002 eliminating the 50-60 Hz noise | Diagram 1 | Drive a 50Hz sinewave Vdc=1.1V Vac=0.3Vpp. Check 20LOG( (amplitude_ecg_codes/gain)/ amplitude_input signal ). Test with no electrode imbalance . Target >80dB - Repeat with 51k¦¦47nF imbalance in one electrode. Target > 40dB | M | CMRR>80dB (no electrode imbalance) CMRR>40dB 51k¦¦47nF electrode imbalance)
37 | ECG | ADPD402 and DG2502 | ADPD4002 input impedance | Performance |  | Enable the 4002 Switch and connec the contacts to the ECG out of interface board | Diagram 1 | Test at 5Hz, 1mVpp. Check output amplitude. Add 1M in series with each input and check change in output amplitude. Calculate Zin | M | >10M
38 | System | DIALOG | Test Bluetooth link | Functional | user feedback | RF section is functional | Diagram 1 | reset dialog, load program (advertising). Check RF link works with USB dongle /phone or other | L | system advertises |  |  | This test needs some examinaion if it could be automated; test signal strength? | Final Functional
|  | DIALOG | Test Bluetooth link | Functional | user feedback | RF section is functional |  | reset dialog, load program (advertising). Check RF link works with USB dobngle /phone or other |  | system advertises
39 | System | DIALOG | UART/SPI interface | Functional | Self Test | Uart loopback test | Diagram 1 | reset dialog, load program (uart loopback). SPI and UART works | L | program downloads and words match |  |  |  | Final Functional
| System | DIALOG | Test Bluetooth link | Functional | user feedback | RF section is functional |  | reset dialog, load program (uart loopback). SPI and UART works |  | system advertises
40 | System | Button | Press button | Functional | user feedback | Push buttons and check response | Diagram 1 | ask user to push button and wait for interrupt | L | interrupt received |  |  |  | Final Functional
41 | Temperature | TEMP. SENSING | Temp Sensing linearity alone | Performance |  | Check circuit linearity simulating the NTC with a Programmable Resistance substitutter or Zbank | Diagram 1 | change the resistor value and obtain the temperature value. Check if the adquired values matches with the expected one | H | tbd |  |  | Use Coolidge for temperature testing?
42 | Temperature | TEMP. SENSING | Temp Sensing Repeatability alone | Performance |  | Check repeatibility simulating the NTC with a Programmable Resistance substitutter or Zbank | Diagram 1 | measure several times the same value and obtain the std deviation | H | tbd |  |  |  | Characterization
43 | Temperature | TEMP. SENSING | Temp Sensing linearity all togheter | Performance |  | Check circuit linearity simulating the NTC with a Programmable Resistance substitutter or Zbanck while the ECG and Z signals are also adquired. Is the noise floor the same? | Diagram 1 | change the resistor value and obtain the value. Check if the adquired values matches with the expected one | L | tbd |  |  |  | Characterization
44 | Temperature | TEMP. SENSING | Temp Sensing Repeatability all togheter | Performance |  | Check repeatibility simulating the NTC with a Programmable Resistance substitutter or Zbanck while the ECG and Z signals are also adquired. Is the noise floor the same? | Diagram 1 | measure several times the same value and obtain the std deviation | L | tbd |  |  |  | Characterization/FF- only few points
45 | AD7156 | I2C communication | Read ID | Functional | Self Test | Read device ID | Diagram 1 | Read the device ID | L | ID is correct |  |  |  | Final Functional
46 | AD7156 | Value out | Use a metal plate to generate an output signal |  |  | Check the output with and without a metal plate placed below the watch | Diagram 1 | TBD | L | Output is observed |  |  |  | Module
48 | ADP5360 | I2C communication | Read ID | Functional | Self Test | Read device ID | Diagram 1 | Read the device ID | L | ID is correct |  |  |  | Final Functional
49 | POWER MANAGEMENT | POWER MANAGEMENT & USB POWER | USB "wall" and "PC" charges battery test | Functional |  | check the USB "Wall" and "PC" supply recharges the battery (USB "PC" and "Wall" connected alternately) | Diagram 1 | mimick battery with keithley and check it sinks current (so USB5V is present and charging the battery) | M | charging current >80mA |  |  | follow up with Dan | Final Functional
50 | POWER MANAGEMENT | POWER MANAGEMENT & USB POWER | USB PC test | Functional |  | check the USB "PC" connection (battery is connected, USB "PC" is connected but USB "Wall" is not connected) | Diagram 1 | set keithley ch1 to 3.8 volts, keithley ch2 is disconnected and ch3 is connected (4.8v) and OTG_SEL means not get voltage from PC. check ch3 current is zero or negative and check USB_ID. Second step: change OTG_SEL, therefore ch3 current will have to be positive, USB_ID will change its level and the USB communication has to work. | M | ADP5350 status is correct |  |  |  | Final Functional
51 | POWER MANAGEMENT | POWER MANAGEMENT & USB POWER | USB Wall test | Functional |  | check the USB "Wall" connection (battery is connected, USB "PC" not connected but USB "Wall" is connected) | Diagram 1 | set keithley ch1 to 3.8 volts, keithley ch2 is connected (5volts) and ch3 is connected (4.8v) and OTG_SEL low. check ch3 current is zero or negative and check USB_ID is floating (set input with pull-up). Second step: change OTG_SEL to high. every thing remains equal. Third step: Ch2 is disconnected, USB_ID goes low and ch3 will supply the FTDI. send message and wait response by USB. | M | ADP5350 status is correct |  |  | 25mV resolution of input. Can possibly utilize AD8233 | Final Functional
52 | POWER MANAGEMENT | POWER MANAGEMENT & USB POWER | Battery test | Functional |  | check supplies. Just the battery is connected to the board (USB "Wall" not connected and USB "PC" not connected) | Diagram 1 | Check if the watch is working without connecting to the Wall or PC charger | L | ADP5350 status is correct |  |  |  | Final Functional
53 | POWER MANAGEMENT |  | Power level test |  |  | Test the power drawn by the board at different power modes |  | TBD | L | TBD |  |  |  | Final Functional
54 | ADP160 | Relibility | Test that LDO produces required volatge when input is changed | Performance |  | Simply toggle input to make sure output is constant | Diagram 1 | Not only can we vary Vin while checking Vout, we can also test the LDO with and without a load. Load a 1k and 100k resistor to Vout to make sure Vout remains the same. |  | equal to Vout ± 5% |  |  |  | Final Functional
Electrodes

## Sheet1
Command | Description
Plot the waveform of the particular test from all the result folders
Get the result value of a particular test from all the result folders
Get the result value of all the tests from result folders
Pick out the test that failed from all the result folders

## Sheet5
Metric ID | What do we care about ? | What metric do we need | Metric defination | Spec | How to interpret? | Failure/Warning
1 | Strength of P,QRS and T waves | SNR metric | Power of the P, QRS and T waves wrt the noise floor of the ECG. | T_SNR > 3dB R_SNR > 20dB | Higher energy = presence of p/qrs/t waves. | Warning for lower SNR values
2 | Strength of P,QRS and T waves | Spectral distribution Ratio (SDR) | Ratio of energy in 5-14 Hz to the energy in 5-50 Hz. | Acceptable : 0.3 < SDR < 0.95 | Very low SNR = presence of high frequency artifacts (muscle movement)Very high SNR = presence of QRS type artifact (electrode movement) | Warning for out of range values
3 | Presence of the correct ECG morphology | Flags to determine the observation of P, QRS, T waves | Visual observation of presence of the P, QRS, T waves | P wave seen? QRS wave seen? T wave seen? | If seen, the correct morphology is observable. | Failure if correct morphology not seen
4 | ECG must not clip or rail | ECG railing | % of time the ECG rails during a capture of xx seconds | Valid ECG data > 80% of each recording for all recordings | Lower railing percentage = better. | Failure if railing percent higher than 80%
Non- ATE (Comprehensive) Test
6 | ECG Signal quality on body | Wear watch on wrist; Make sure AD5940 and ADPD4002 Switches are off; Turn on the 8233 Switch and collect AD5940 output data | a. Presence of the correct ECG morphology; b. Strength of P,QRS and T waves; c. ECG does not clip or rail | a. P wave seen? QRS wave seen? T wave seen?; b. T_SNR > 3dB R_SNR > 20dB and 0.3 < SDR < 0.95 c.Valid ECG data > 80% of each recording for all recordings | None | Few
7 | ECG Signal quality with ECG simulator in relation with PPG, accelerometer and temperature (use case #4) | Wear the watch and Run ECG@1000Hz, PPG@ 50Hz, ADXL@50Hz and temp@1Hz | 1. ECG morphology; 2. ECG SNR; 3. PPG morphology; 4. PPG SNR (MI); 5. Data integrity | 1. ECG morphology is as per requirement; 2. ECG SNR is not reduced; 3. PPG morphology is as per requirement; 4. PPG MI is not reduced; 5. No packet loss | None | Few
8 | PPG Moduation Index | Feed PPG signal with known modulation index using PPG simulator to Watch | PPG Modulation index value | Match the Modulation index to received signal | None | Few
9 | PPG Test with AGC in relation to adxl and temp (use case #1) | Wear the watch and Run PPG@ 500Hz, ADXL@50Hz and temp@1Hz | 1. PPG morphology; 2. PPG SNR; 3. Data integrity | 1. PPG morphology as per requirement; 2. PPG MI does not decrease; 3. no data loss | None | Few
10 | PPG Test with AGC in relation to EDA, adxl and temp (use case #2) | Wear the watch and Run EDA@30Hz, PPG@ 100Hz, ADXL@50Hz and temp@1Hz | 1. PPG morphology; 2. EDA value is correct;3 PPG SNR (MI); 4. Data integrity | 1. PPG morphology as per requirement; 2. EDA value within range; 3. PPG MI does not decrease; 4. no data loss | None | Few
11 | PPG Test with AGC in relation to ECG, adxl and temp (use case #3) | Run ECG@250Hz, PPG@ 100Hz, ADXL@50Hz and temp@1Hz | 1. PPG morphology; 2. ECG morphology;3 PPG SNR (MI); ECG SNR; 5. Data integrity | 1. PPG morphology as per requirement; 2. ECG morphology as per requirement; 3. PPG MI does not decrease; 4. ECG SNR does not decrease; 5. no data loss | None | Few