## qa_test_plan
Test | Pre-requisites | Test methodology | Pass/Fail creteria | Availability | Priority | Test type | Comments/Reference
General Tests
FW Download Test |  | Connect to the Watch using USB, enter bootloader mode and download the firmware | Process completes without any errors | Implemented
USB Connection Test |  | Able to connect to the watch using USB and read firmware version | connects without any error and returns firmware version | Implemented
BLE Connection Test |  | Able to connect to the watch using BLE and read firmware version | connects without any error and returns firmware version | Implemented
BLE Throughput |  | Check if expected data rate is reachable, for required time at required distance between Dongle and Watch | Need to run the test for 1 hour, at 10m distance and achieve max throughput
USB Throughput |  | Check if expected data rate is reachable, for required time using the USB cable between ATE test script and Watch | Need to run the test for 1 hour and achieve max throughput | Implemented |  |  | Being tested as part of Usecase tests
Date and Time Test |  | Set the Date and time to a known date and time and read back the Date and Time | Read back Date and Time matches the written date and time | Implemented
Get System Info |  | Fetch system info related to current Watch being tested, which will be captured in the ATE log report for the unit being tested | Should be able to read correct MAC Address of the device
ADPD4xxx
Get Device ID |  | read the device id and verify the id | Device id matches the expected id | Implemented
DCB Test |  | write the dcb file with ADPD configuration and read back registers and compare values | 1. DCB Write data matches with DCB Read data 2. DCB Write data matches with Register Values | Implemented
DCB ADPD Stream Test |  | write the dcb file with ADPD configuration and read back registers and compare values | 1. DCB Write data matches with DCB Read data 2. DCB Write data matches with Register Values 3. Stream data without sample loss and correct freq | Implemented
DCB Repeatability Test |  | Perform the DCB Test and DCB Stream test for 10 iterations | No failures to be found on either of the tests in any of the iterations | Implemented
ADPD Stream Test |  | Stream data using Green, Red, IR and Blue LEDs (Slot F, G, H and I) and verify the stream data | Log file has data in the correct format and as per the sweep sampling frequency | Implemented
ADPD Stream NAND Flash Logging Test |  | Enable NAND Flash logging and start stream for 10 seconds and stop stream. Download the log file and vefify data. | Log file has data in the correct format and as per the sweep sampling frequency | Implemented |  |  | convert m2m2 to csv - pending
Frequency Sweep ADPD Stream Test |  | Sweep across valid frequencies (50, 100, 500) and perform stream test | Log file has data in the correct format and as per the sweep sampling frequency | Implemented
Frequency Sweep Repeatability ADPD Stream Test |  | Sweep across valid frequencies (50, 100, 500) and perform stream test for 10 iterations | Log file has data in the correct format and as per the sweep sampling frequency | Implemented
ADPD AGC Test |  | Turn ON AGC and collect data and Turn OFF AGC and collect data and compare the data for each LEDs(Green, Red, Blue and IR) | AGC ON data values are higher than AGC OFF data values | Implemented
Frequency Sweep ADPD AGC Test |  | Perform the ADPD AGC test across 50, 100 and 500 Hz frequencies | AGC ON data values are higher than AGC OFF data values | Implemented
PPG Stream Test |  | Start Stream, collect logs for 30 seconds, Stop Stream | Log file has data in the correct format and as per the default sampling frequency | Implemented
PPG Stream NAND Flash Logging Test |  | Enable NAND Flash logging and start stream for 10 seconds and stop stream. Download the log file and vefify data. | Log file has data in the correct format and as per the sweep sampling frequency |  |  |  | convert m2m2 to csv - pending
PPG AGC Test |  | Turn ON AGC and collect data and Turn OFF AGC and collect data and compare the data | AGC ON data values are higher than AGC OFF data values | InReview
ADXL
Get Device ID |  | read the device id and verify the id | Device id matches the expected id | Implemented
DCB Test |  | write the dcb file with ADXL configuration and read back registers and compare values | Values in the DCB file and read back register values match | Implemented
DCB Stream Test |  | write the dcb file with ADXL configuration and Start Stream, collect logs for 30 seconds, Stop Stream | Log file has data in the correct format and as per the default sampling frequency | InReview
DCB Repeatability Test |  | Perform the DCB Test and DCB Stream test for 10 iterations | No failures to be found on either of the tests in any of the iterations | Implemented
ADXL Stream Test |  | Start Stream, collect logs for 30 seconds, Stop Stream | Log file has data in the correct format and as per the default sampling frequency | Implemented
ADXL Stream NAND Flash Logging Test |  | Enable NAND Flash logging and start stream for 10 seconds and stop stream. Download the log file and vefify data. | Log file has data in the correct format and as per the sweep sampling frequency |  |  |  | convert m2m2 to csv - pending
Frequency Sweep Stream Test |  | Sweep across valid frequencies and perform stream test | Log file has data in the correct format and as per the sweep sampling frequency | InReview
Frequency Sweep Repeatability Stream Test |  | Sweep across valid frequencies and perform stream test with 5 iterations | Log file has data in the correct format and as per the sweep sampling frequency | InReview
AD5940
Get Device ID |  | read the device id and verify the id | Device id matches the expected id | Implemented
ECG DCB Test |  | write the dcb file with ADPD configuration and read back registers and compare values | Values in the DCB file and read back register values match | InReview
ECG DCB Stream Test |  | write the dcb file with ADPD configuration and read back registers and compare values | 1. DCB Write data matches with DCB Read data 2. DCB Write data matches with Register Values 3. Stream data without sample loss and correct freq | InReview
ECG DCB Repeatability Test |  | Perform the DCB Test and DCB Stream test for 10 iterations | No failures to be found on either of the tests in any of the iterations | InReview
ECG Stream Test |  | Start Stream, collect logs for 30 seconds, Stop Stream | Log file has data in the correct format and as per the default sampling frequency | Implemented
ECG Stream NAND Flash Logging Test |  | Enable NAND Flash logging and start stream for 10 seconds and stop stream. Download the log file and vefify data. | Log file has data in the correct format and as per the sweep sampling frequency |  |  |  | convert m2m2 to csv - pending
AD5940 Switch Test |  | Collect stream data with Switch Turned ON and OFF and compare data | No data available in Switch OFF stream file and valid data in Siwtch ON stream file | Implemented
Frequency Sweep ECG Stream Test |  | Sweep across valid frequencies and perform stream test | Log file has data in the correct format and as per the sweep sampling frequency | InReview
Repeatability Frequency Sweep ECG Stream Test |  | Sweep across valid frequencies and perform stream test with 5 iterations | Log file has data in the correct format and as per the sweep sampling frequency | InReview
EDA DCB Test |  | write the dcb file with ADPD configuration and read back registers and compare values | Values in the DCB file and read back register values match | InReview
EDA DCB Stream Test |  | write the dcb file with ADPD configuration and read back registers and compare values | 1. DCB Write data matches with DCB Read data 2. DCB Write data matches with Register Values 3. Stream data without sample loss and correct freq | InReview
EDA DCB Repeatability Test |  | Perform the DCB Test and DCB Stream test for 10 iterations | No failures to be found on either of the tests in any of the iterations | InReview
EDA Stream Test |  | Start Stream, collect logs for 30 seconds, Stop Stream | Log file has data in the correct format and as per the default sampling frequency | Implemented
EDA Stream NAND Flash Logging Test |  | Enable NAND Flash logging and start stream for 10 seconds and stop stream. Download the log file and vefify data. | Log file has data in the correct format and as per the sweep sampling frequency |  |  |  | convert m2m2 to csv - pending
Frequency Sweep EDA Stream Test |  | Sweep across valid frequencies and perform stream test | Log file has data in the correct format and as per the sweep sampling frequency | InReview
Repeatability Frequency EDA Sweep Stream Test |  | Sweep across valid frequencies and perform stream test with 5 iterations | Log file has data in the correct format and as per the sweep sampling frequency | InReview
ADP5360
Get Device ID |  | read the device id and verify the id | Device id matches the expected id | Implemented
Get the battery voltage |  | Issue command to get the battery voltage, to compare with the real voltage. | getBatteryInfo Value read from test application should match with measured value.
AD7156
Get Device ID |  | read the device id and verify the id | Device id matches the expected id | Implemented
DCB Test |  | write the dcb file with ADPD configuration and read back registers and compare values | Values in the DCB file and read back register values match
DCB Low Touch Stream Test |  | write the dcb file with ADPD configuration and read back registers and compare values | 1. DCB Write data matches with DCB Read data 2. DCB Write data matches with Register Values 3. Stream data without sample loss and correct freq |  |  |  | ToDo From Firmware
DCB Repeatability Test |  | Perform the DCB Test and DCB Stream test for 10 iterations | No failures to be found on either of the tests in any of the iterations
Check touch response |  | Issue a command to the watch,to check the touch response | Command returns valid output values
Temperature
Get Device ID |  | read the device id and verify the id | Device id matches the expected id |  |  |  | Same as ADPD
DCB Test |  | write the dcb file with ADPD configuration and read back registers and compare values | Values in the DCB file and read back register values match
Stream Test |  | Start Stream, collect logs for 30 seconds, Stop Stream | Log file has data in the correct format and as per the default sampling frequency | Implemented
Stream NAND Flash Logging Test |  | Enable NAND Flash logging and start stream for 10 seconds and stop stream. Download the log file and vefify data. | Log file has data in the correct format and as per the sweep sampling frequency |  |  |  | convert m2m2 to csv - pending
Frequency Sweep Stream Test |  | Sweep across valid frequencies and perform stream test | Log file has data in the correct format and as per the sweep sampling frequency |  |  |  | Not needed
Repeatability Frequency Sweep Stream Test |  | Sweep across valid frequencies and perform stream test with 5 iterations | Log file has data in the correct format and as per the sweep sampling frequency |  |  |  | Not needed
File System
Read Flash ID |  | Issue a command to get the Flash ID. | Readback 0x2000 | Implemented
Format the file system |  | Format the Flash and check flash status | Flash status OK | Implemented |  |  | part of FS Stream test
Volume Info |  | Check the total space and used space in flash with known data size in flash | Volume Info returns valid values for total space and space in flash
Directory Listing |  | Run the test to create files beyond 63 and see if it fails. Get the directory listing to verify maximum number of files present is 63
File Download |  | Create a file with known pattern. Download and verify if its properly downloaded | Downloaded files pattern matches the known pattern | Implemented |  |  | part of FS Stream test
ECC test |  | Enabling and disabling ECC and doing tests |  |  |  |  | ToDo From Firmware
Low Battery Logging |  | Fimware adds a limit on logging based on battery level( 5%  low, 1% critical). The test is : drain the battery, start logging, check if at critical level logs are properly saved or not? |  |  |  |  | ToDo From Firmware
Bad Block Testing
Misc Tests
DCB Tests |  | Perform ADPD4000, ADXL, ECG, PPG, EDA and AD7150 DCB tests one after other and check for failures | All tests passes
DCB Repeatability Tests |  | Perform ADPD4000, ADXL, ECG, PPG, EDA and AD7150 DCB tests one after other for 10 iterations and check for failures | All tests passes
ECG Switch Independence Test |  | Toggle ECG switch OFF and run ADPD, ADXL, EDA and Temp streams and verify that the ECG switch is not affecting other streams | All streams log valid data when irrespective of the ECG Switch state
Usecase Tests
Usecase-1 Test |  | Run all the streams which are part of the usecase and verify packet losses on all streams and verify NAND Flash Logging | Log file has data in the correct format and as per the sampling frequency. NAND Flash log file has data in correct format and as per sampling freq | InReview
Usecase-1 Radomized Stream Sequence Test |  | Exercise random order or all combination of sequential order for different stream starts and stream stop commands in usecase tests | All usecases passes irrespective of the order of stream commands | InReview
Usecase-2 Test |  | Run all the streams which are part of the usecase and verify packet losses on all streams | Log file has data in the correct format and as per the sampling frequency. NAND Flash log file has data in correct format and as per sampling freq | InReview
Usecase-2 Radomized Stream Sequence Test |  | Exercise random order or all combination of sequential order for different stream starts and stream stop commands in usecase tests | All usecases passes irrespective of the order of stream commands | InReview
Usecase-3 Test |  | Run all the streams which are part of the usecase and verify packet losses on all streams | Log file has data in the correct format and as per the sampling frequency. NAND Flash log file has data in correct format and as per sampling freq | InReview
Usecase-3 Radomized Stream Sequence Test |  | Exercise random order or all combination of sequential order for different stream starts and stream stop commands in usecase tests | All usecases passes irrespective of the order of stream commands | InReview
Usecase-4 Test |  | Run all the streams which are part of the usecase and verify packet losses on all streams | Log file has data in the correct format and as per the sampling frequency. NAND Flash log file has data in correct format and as per sampling freq | InReview
Usecase-4 Radomized Stream Sequence Test |  | Exercise random order or all combination of sequential order for different stream starts and stream stop commands in usecase tests | All usecases passes irrespective of the order of stream commands | InReview
Usecase-5 Test |  | Run all the streams which are part of the usecase and verify packet losses on all streams | Log file has data in the correct format and as per the sampling frequency. NAND Flash log file has data in correct format and as per sampling freq | InReview
Usecase-5 Radomized Stream Sequence Test |  | Exercise random order or all combination of sequential order for different stream starts and stream stop commands in usecase tests | All usecases passes irrespective of the order of stream commands | InReview
Randomized all usecases Test |  | Exercise random order or all combination of sequential order for running 5 usecases | All usecases passes irrespective of the order | InReview

## Feedback
Sep 28, 2020 (Dan)
BLE Throughput | DA: Why not go for a higher data rate? I thought ~6.5 kB/s was the target | We should go for higher data rate. I'll implement it at max supported rate.
USB Throughput | DA: Same as above (but higher) | Same as above
Date and Time Test | DA: Is it worth checking whther this persists correctly after a power mode / reset? | It will be a good test case. I'll include this in the test plan. The reset command in CLI has some issue which needs to be fixed to implement this though. I'll figure out a way to do it.
Get System Info | DA: If possible, I think it would be great to get a full readout of everything available. MAC addresses, configuration registers, ChipID's….. All of it. If/When we implement a Valencia style test it would be a great way to ensure the device is going into the right state | I'll modify this test case to cover these aspects and share the details. We can collect the configuration registers values and compare it against a known set of values after Reset
DCB Test | DA: Does this mean that you trigger the watch to load the DCB before you check? | Yes, as part of this DCB test, we write a known DCB into watch and trigger the watch to load the DCB. Afte that we read the registers and makes sure the DCB write and load functionality worked. Checks are implemented in 2 stages, 1. Reading back DCB and verifying values to ensure write DCB worked 2. Reading back registers and verifying values to ensure load DCB worked
ADPD Stream Test | DA: Stream over what… USB? I would also pick the highest supported frequency | Currently the whole test will use either USB/BLE connectivity and this option will be selected in the initialization of the test. The entire suite can be run for both USB and BLE if that sounds fine.
Frequency Sweep ADPD Stream Test | DA: All LEDs individually? Only one? | This will cover all LEDS, first individually one after other and later all LEDs at the same time.
ADPD AGC Test | DA: I think this could fail depending on the way the watch is setup (mechanically)… worth keeping an eye on | Noted.
PPG Stream Test | DA: Is PPG meaningful still if we're using Use Cases? How is this any different than Use Case 1? | This test case will be same as Usecase 1. I'll remove this entry. But this will still be added as a function in order to be reused if needed.
Get the battery voltage | DA: It would be helpful if you sandwiched this test at the beginning and end of a full test sequence. Monitoring that begin/end delta would help uncover if we put in some change that pulls substantially more power. (Note: You would need to somehow disable charging to do this....not sure if that's easy to do) | This sounds like a good way of testing it. I'll check how to implement this and modify the test case.
Low Battery Logging | DA: I wouldn't advocate you do this by actually draining the battery. You should add a test command that mimics that response from the battery so you don't have to wait that full time | I'll check with the firmware team if we can have a command to mimic the response for testing this. I'll update the testcase accordingly.
| DA: Can we have a test that monitors total # of bad blocks? If we use one watch all the time we should see this slowly increase | Sure we can have this test case. I'll not be checking the Failure criteria for this testcase as it is expected to have some bad blocks as we test more. We can have it as a monitoring test case as you mentioned.