## TopPage
Command Summary:
Getting Version
GetVersion
Setting Date and Time
SetDateTime
ADPD register read/write and PPG application library configuration read/write
| Register Read or Write |  |  | PPG Configuration Setting Read or Write
ADPD stream
| ADPD Start Commands |  |  | ADPD Stop Commands |  |  | ADPD Data
ADXL Stream
| ADXL Start Commands |  |  | ADXL Stop Commands |  |  | ADXL Data
PPG (HR) and sychronized ADPD-ADXL Stream
| PPG Start Commands |  |  | PPG Stop Commands |  |  | PPG Data
ADPD ADXL Sync Data
EDA Stream
| EDA Start Commands |  |  | EDA Stop Commands |  |  | EDA Data
Temperature Stream
| Temperature Start Commands |  |  | Temperature Stop Commands |  |  | Temperature Data
| PPG Motion Compensated Stream Start Commands |  |  | PPG Motion Compensated Stream Stop Commands |  |  | PPG Motion Compensated Data
Low Touch
Low touch commands to activate / deactivate the cap sensors for on/off wrist event detection
ECG Stream
| ECG start | ECG Stop |  | ECG Data
Pedometer Stream
| Pedometer Start | Pedometer Stop |  | Pedometer Data
Signal Quality Index for PPG signal
| SQIStart | SQIStop |  | SQIData
Logging of data to Nand Flash
| File System Logging commands |  | PPG Logging Start |  | PPG Logging Stop | ECG Logging Start |  | ECG Logging Stop

## GetVersion
Get Version Command | [9 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination |  | M2M2_ADDR_SYS_PM = 0xC500
[10 empty] | Request Command |  | M2M2_APP_COMMON_CMD_GET_VERSION_REQ = 0x0 |  |  |  |  | Command to get the version of Mother board
PC to Device | [9 empty] | Response Command |  | M2M2_APP_COMMON_CMD_GET_VERSION_RESP = 0x1
[10 empty] | Status |  | M2M2_APP_COMMON_STATUS_OK = 0x0 |  |  |  |  | Version details obtained successfully
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C7 | 05 | C5 | 0 | 00 | 42 | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24
Description | Command | Status | Major version LSB | Major version MSB | Minor version LSB | Minor version MSB | Patch version LSB | Pacth version MSB | version string ASCII byte | version string ASCII byte | version string ASCII byte | version string ASCII byte | version string ASCII byte | version string ASCII byte | version string ASCII byte | version string ASCII byte
Data | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00
Payload
Byte Number | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40
Description | version string ASCII byte | version string ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte
Data | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00
Payload
Byte Number | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56
Description | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte
Data | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00
Payload
Byte Number | 57 | 58 | 59 | 60 | 61 | 62 | 63 | 64 | 65 | 66
Description | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte
Data | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00
Device to PC
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C5 | 80 | C7 | 05 | 00 | 42 | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24
Description | Command | Status | Major version LSB | Major version MSB | Minor version LSB | Minor version MSB | Patch version LSB | Pacth version MSB | version string ASCII byte | version string ASCII byte | version string ASCII byte | version string ASCII byte | version string ASCII byte | version string ASCII byte | version string ASCII byte | version string ASCII byte
Data | 01 | 00 | 03 | 00 | 05 | 00 | 00 | 00
Payload
Byte Number | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40
Description | version string ASCII byte | version string ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte
Data
Payload
Byte Number | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56
Description | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte
Data
Payload
Byte Number | 57 | 58 | 59 | 60 | 61 | 62 | 63 | 64 | 65 | 66
Description | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte | Build version ASCII byte
Data

## HibernateModeControl
get_hib_mode_status | [8 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[9 empty] | Destination | M2M2_ADDR_SYS_PM = 0xC500
[9 empty] | Request Command | M2M2_PM_SYS_GET_HIBERNATE_MODE_STATUS_REQ = 0x9A | To get the hibernate mode status- enabled/disabled in the Watch Fw. It is based on this value that Hibernate Mode happens. Hibernate Mode or Auto Power off of the Watch happens after a timeout of 5 mins, if there is no BLE/USB connection, No logging in progress, No button keys are pressed, No streaming happening on Watch display
PC to Device | [8 empty] | Response Command | M2M2_PM_SYS_GET_HIBERNATE_MODE_STATUS_RESP = 0x9B
[9 empty] | Status | M2M2_PM_SYS_STATUS_OK = 0x41 | "Hib_mode_status" details obtained successfully
Header | [8 empty] | Hib_mode_status | 1 -> Hibernate Mode enabled 0-> Hibernate Mode disabled
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C7 | 05 | C5 | 0 | 00 | 0B | 00 | 00
Payload
Byte Number | 9 | 10 | 11
Description | Command | Status | Hib_mode_status
Data | 9A | 41 | 00
Device to PC
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C5 | 0 | C7 | 05 | 00 | 0B | 00 | 00
Payload
Byte Number | 9 | 10 | 11
Description | Command | Status | Hib_mode_status
Data | 9B | 41 | 01
set_hib_mode_status | [8 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[9 empty] | Destination | M2M2_ADDR_SYS_PM = 0xC500
[9 empty] | Request Command | M2M2_PM_SYS_SET_HIBERNATE_MODE_STATUS_REQ = 0x9C | To set the hibernate mode status- enabled/disabled in the Watch Fw. It is based on this value that Hibernate Mode happens. Hibernate Mode or Auto Power off of the Watch happens after a timeout of 5 mins, if there is no BLE/USB connection, No logging in progress, No button keys are pressed, No streaming happening on Watch display
PC to Device | [8 empty] | Response Command | M2M2_PM_SYS_SET_HIBERNATE_MODE_STATUS_RESP = 0x9D
[9 empty] | Status | M2M2_PM_SYS_STATUS_OK = 0x41 | "Hib_mode_status" changed successfully
Header | [8 empty] | Hib_mode_status | 1 -> Hibernate Mode enabled 0-> Hibernate Mode disabled
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C7 | 05 | C5 | 0 | 00 | 0B | 00 | 00
Payload
Byte Number | 9 | 10 | 11
Description | Command | Status | Hib_mode_status
Data | 9C | 41 | 00
Device to PC
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C5 | 0 | C7 | 05 | 00 | 0B | 00 | 00
Payload
Byte Number | 9 | 10 | 11
Description | Command | Status | Hib_mode_status
Data | 9D | 41 | 00

## Max_Tx_pkt_comb_Cnt
get_max_tx_pkt_comb_cnt | [8 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[9 empty] | Destination | M2M2_ADDR_SYS_PM = 0xC500
[9 empty] | Request Command | M2M2_PM_SYS_BLE_GET_MAX_TX_PKT_COMB_CNT_REQ = 0x92 | get the "max_tx_pkt_comb_cnt" used in the Watch Fw. It is based on this value that no: of pkts for BLE Tx is combined and send out.
PC to Device | [8 empty] | Response Command | M2M2_PM_SYS_BLE_GET_MAX_TX_PKT_COMB_CNT_RESP = 0x93
[9 empty] | Status | M2M2_PM_SYS_STATUS_OK = 0x41 | "max_tx_pkt_comb_cnt" details obtained successfully
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C7 | 05 | C5 | 0 | 00 | 0B | 00 | 00
Payload
Byte Number | 9 | 10 | 11
Description | Command | Status | max_tx_pkt_comb_cnt
Data | 92 | 41 | 00
Device to PC
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C5 | 0 | C7 | 05 | 00 | 0B | 00 | 00
Payload
Byte Number | 9 | 10 | 11
Description | Command | Status | max_tx_pkt_comb_cnt
Data | 93 | 41 | 04
set_max_tx_pkt_comb_cnt | [8 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[9 empty] | Destination | M2M2_ADDR_SYS_PM = 0xC500
[9 empty] | Request Command | M2M2_PM_SYS_BLE_SET_MAX_TX_PKT_COMB_CNT_REQ = 0x94 | set the "max_tx_pkt_comb_cnt" to be used in the Watch Fw. It is based on this value that no: of pkts for BLE Tx is combined and send out.
PC to Device | [8 empty] | Response Command | M2M2_PM_SYS_BLE_SET_MAX_TX_PKT_COMB_CNT_RESP = 0x95
[9 empty] | Status | M2M2_PM_SYS_STATUS_OK = 0x41 | "max_tx_pkt_comb_cnt" details that was set successfully
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C7 | 05 | C5 | 0 | 00 | 0B | 00 | 00
Payload
Byte Number | 9 | 10 | 11
Description | Command | Status | max_tx_pkt_comb_cnt
Data | 94 | 41 | 02
Device to PC
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C5 | 0 | C7 | 05 | 00 | 0B | 00 | 00
Payload
Byte Number | 9 | 10 | 11
Description | Command | Status | max_tx_pkt_comb_cnt
Data | 95 | 41 | 02

## SetDateTime
[12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
Set Date and Time | [11 empty] | Destination |  | M2M2_ADDR_SYS_PM = 0xC500
[12 empty] | Request Command |  | M2M2_PM_SYS_COMMAND_SET_DATE_TIME_REQ = 0x42 | Command to set current Date and Time on device
[12 empty] | Response Command |  | M2M2_PM_SYS_COMMAND_SET_DATE_TIME_RESP = 0x43
PC to Device | [11 empty] | Status |  | M2M2_APP_COMMON_STATUS_OK =0x 0 (or) | Date and Time set successfully
[14 empty] | M2M2_PM_SYS_STATUS_ERR_ARGS = 0x42 (or) | Error in Date and Time format
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C7 | 05 | C5 | 00 | 00 | 15 | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
Description | Command | Status | Year LSB | year MSB | month | day | hour | minute | second | TimeZone LSB | TimeZone | TimeZone | TimeZone MSB
Data | 42 | 00 | E5 | 07 | 01 | 12 | 0E | 3A | 1C | 58 | 4D | 00 | 00
Device to PC
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C5 | 00 | C7 | 05 | 00 | 0A | 00 | 00
Payload
Byte Number | 9 | 10
Description | Command | Status
Data | 43 | 00

## LT Commands
Low Touch Commands
[13 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
| Set Date and Time | [11 empty] | Destination |  | M2M2_ADDR_SYS_PM = 0xC500
[13 empty] | Request Command |  | M2M2_PM_SYS_COMMAND_SET_DATE_TIME_REQ = 0x42 |  |  | Command to set current Date and Time on device
[13 empty] | Response Command |  | M2M2_PM_SYS_COMMAND_SET_DATE_TIME_RESP = 0x43
| PC to Device | [11 empty] | Status |  | M2M2_APP_COMMON_STATUS_OK =0x 0 (or) |  |  | Date and Time set successfully
[15 empty] | M2M2_PM_SYS_STATUS_ERR_ARGS = 0x42 (or) |  |  | Error in Date and Time format
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C5 | 00 | 00 | 15 | 00 | 00 | [7 empty] | Ignore this command if Date and Time are already set
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
| Description | Command | Status | Year LSB | year MSB | month | day | hour | minute | second | TimeZone LSB | TimeZone | TimeZone | TimeZone MSB
| Data | 42 | 00 | E5 | 07 | 01 | 12 | 0E | 3A | 1C | 58 | 4D | 00 | 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 00 | C7 | 05 | 00 | 0A | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10
| Description | 1: Command; 2: Status
| Data | 1: 43; 2: 00
[13 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
| Activate Cap Sensor Command | [11 empty] | Destination |  | M2M2_ADDR_SYS_PM = 0xC500
[13 empty] | Request Command |  | M2M2_PS_SYS_COMMAND_ACTIVATE_TOUCH_SENSOR_REQ = 0x4A |  |  | Command to activate cap sensors on PM board
[13 empty] | Response Command |  | M2M2_PS_SYS_COMMAND_ACTIVATE_TOUCH_SENSOR_RESP = 0x4B
| PC to Device | [11 empty] | Status |  | M2M2_PM_SYS_STATUS_OK =0x41 |  |  | Cap sensors activated successfully
[15 empty] | M2M2_PM_SYS_STATUS_ERR_ARGS = 0x42 (or) |  |  | Error in activate touch sensor command
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C5 | 00 | 00 | 0A | 00 | 00
| Payload | [11 empty] | After this command, the cap sensors are active to monitor on/off wrist event.
| Byte Number | 1: 9; 2: 10
| Description | 1: Command; 2: Status |  |  |  | On Wrist: Once the watch is placed firmly on the wrist for more than 5 seconds consistently,
| Data | 1: 86; 2: 00 |  |  |  | On wrist event message will be sent out, if LT config file is present
| Device to PC | [11 empty] | Off Wrist: Once the watch is taken off from the wrist for more than 3 seconds, Off wrist
event message will be sent out, if LT config file is present.
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 00 | C7 | 05 | 00 | 0A | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10
| Description | 1: Command; 2: Status
| Data | 1: 87; 2: 41
ON WIRST / OFF WRIST EVENT MESSAGE PACKETS with NAND config file as LT Config file
[14 empty] | Source |  | M2M2_ADDR_SYS_PM = 0xC500
[14 empty] | Destination |  | M2M2_ADDR_APP_CLI = 0xC705
| ON WRIST EVENT MESSAGE | [12 empty] | Command |  | M2M2_PM_SYS_COMMAND_ENABLE_USER_CONFIG_LOG_RESP = 0x61 |  |  | On wrist event is detected, LT cfg based logging enabled | bar: 1 to 2
[14 empty] | Status |  | M2M2_PM_SYS_STATUS_USER_CONFIG_LOG_ENABLED =0x47 |  |  | status ok
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 00 | C7 | 05 | 00 | 0A | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10
| Description | 1: Command; 2: Status
| Data | 1: 60; 2: 00
[14 empty] | Source |  | M2M2_ADDR_SYS_PM = 0xC500
| OFF WRIST EVENT MESSAGE | [12 empty] | Destination |  | M2M2_ADDR_APP_CLI = 0xC705 | bar: 1 to 2
[14 empty] | Command |  | M2M2_PM_SYS_COMMAND_DISABLE_USER_CONFIG_LOG_RESP = 0x63 |  |  | Off wrist event is detected, LT cfg based logging disabled
| Device to PC | [12 empty] | Status |  | M2M2_PM_SYS_STATUS_USER_CONFIG_LOG_DISABLED =0x49 |  |  | status ok
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 00 | C7 | 05 | 00 | 0A | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10
| Description | 1: Command; 2: Status
| Data | 1: 62; 2: 00
ON WIRST / OFF WRIST EVENT MESSAGE PACKETS with gen_blk_DCB as LT Config file
[14 empty] | Source |  | M2M2_ADDR_SYS_PM = 0xC500
[14 empty] | Destination |  | M2M2_ADDR_APP_CLI = 0xC705
| ON WRIST EVENT MESSAGE | [12 empty] | Command |  | M2M2_PM_SYS_COMMAND_ENABLE_USER_CONFIG_LOG_RESP = 0x61 |  |  | On wrist event is detected, LT cfg based logging enabled | bar: 1 to 2
[14 empty] | Status |  | M2M2_PM_SYS_STATUS_DCB_CONFIG_LOG_ENABLED =0x51 |  |  | status ok
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source LSB | Source MSB | Destination LSB | Destination MSB | Length LSB | Length MSB | Checksum LSB | Checksum MSB
| Data | C5 | 00 | C7 | 05 | 00 | 0A | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10
| Description | 1: Command; 2: Status
| Data | 1: 60; 2: 00
[14 empty] | Source |  | M2M2_ADDR_SYS_PM = 0xC500
| OFF WRIST EVENT MESSAGE | [12 empty] | Destination |  | M2M2_ADDR_APP_CLI = 0xC705 | bar: 1 to 2
[14 empty] | Command |  | M2M2_PM_SYS_COMMAND_DISABLE_USER_CONFIG_LOG_RESP = 0x63 |  |  | Off wrist event is detected, LT cfg based logging disabled
| Device to PC | [12 empty] | Status |  | M2M2_PM_SYS_STATUS_DCB_CONFIG_LOG_DISABLED =0x53 |  |  | status ok
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 00 | C7 | 05 | 00 | 0A | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10
| Description | 1: Command; 2: Status
| Data | 1: 62; 2: 00
[13 empty] | Source |  | M2M2_ADDR_APP_CLI = 0xC705
| Deactivate Cap Sensor Command | [11 empty] | Destination |  | M2M2_ADDR_SYS_PM = 0xC500
[13 empty] | Request Command |  | M2M2_PS_SYS_COMMAND_DEACTIVATE_TOUCH_SENSOR_REQ = 0x4C |  |  | Command to deactivate cap sensors on PS board
[13 empty] | Response Command |  | M2M2_PS_SYS_COMMAND_DEACTIVATE_TOUCH_SENSOR_RESP = 0x4D
| PC to Device | [11 empty] | Status |  | M2M2_PM_SYS_STATUS_OK =0x41 |  |  | Cap sensors deactivated successfully
[15 empty] | M2M2_PM_SYS_STATUS_ERR_ARGS = 0x42 (or) |  |  | Error in deactivate touch sensor command
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C5 | 00 | 00 | 0A | 00 | 00 |  |  |  | After this command, the cap sensors are no longer active to monitor on/off wrist event.
Payload
| Byte Number | 1: 9; 2: 10
| Description | 1: Command; 2: Status
| Data | 1: 88; 2: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 00 | C7 | 05 | 00 | 0A | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10
| Description | 1: Command; 2: Status
| Data | 1: 89; 2: 41
[14 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
| get_low_touch_status | [12 empty] | Destination |  | M2M2_ADDR_SYS_PM = 0xC500
[14 empty] | Request Command |  | M2M2_PS_SYS_COMMAND_GET_LOW_TOUCH_LOGGING_STATUS_REQ = 0x66
[14 empty] | Response Command |  | M2M2_PS_SYS_COMMAND_GET_LOW_TOUCH_LOGGING_STATUS_RESP = 0x67 |  |  | Command to low touch application status
|  | 1: PC to Device |  |  |  |  | Status |  | M2M2_PM_SYS_STATUS_LOW_TOUCH_LOGGING_IN_PROGRESS = 0x4B |  |  | LT logging is in progress
[16 empty] | M2M2_PM_SYS_STATUS_LOW_TOUCH_LOGGING_NOT_STARTED = 0x4C |  |  | LT logging not started yet
Header
|  | Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
|  | Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
|  | Data | C7 | 05 | C5 | 00 | 00 | 0A | 00 | 00
Payload
|  | Byte Number | 1: 9; 2: 10
|  | Description | 1: Command; 2: Status
|  | Data | 1: 66; 2: 00
Device to PC
Header
|  | Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
|  | Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
|  | Data | C5 | 00 | C7 | 05 | 00 | 0A | 00 | 00
Payload
|  | Byte Number | 1: 9; 2: 10
|  | Description | 1: Command; 2: Status
|  | Data | 1: 67; 2: 4B

## PpgStart
PPG Start Commands
Load Adpd configuration | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_SENSOR_ADPD4000 = 0xC110
[13 empty] | Request Command |  | M2M2_SENSOR_ADPD_COMMAND_LOAD_CFG_REQ = 0x42 | Command to Load Adpd device configuration
| PC to Device | [11 empty] | Response Command |  | M2M2_SENSOR_ADPD_COMMAND_LOAD_CFG_RESP = 0x43
[13 empty] | Device ID |  | M2M2_SENSOR_ADPD4000_DEVICE_4000_G = 0x28
| Header | [11 empty] | Status |  | M2M2_APP_COMMON_STATUS_OK = 0 (or) | ADPD device loaded successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | [5 empty] | M2M2_APP_COMMON_STATUS_ERROR = 1 (or) | Error in Loading ADPD device
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 0 | 15 | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
| Description | Command | Status | Device ID | return data | return data | return data | return data | return data | return data | return data | return data | return data | return data
| Data | 42 | 00 | 28 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 10 | C7 | 05 | 00 | 15 | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
| Description | Command | Status | Device ID | return data | return data | return data | return data | return data | return data | return data | return data | return data | return data
| Data | 43 | 00 | 28 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00
Clock Calibration |  | 20: source | M2M2_ADDR_APP_CLI = 0xC705
|  | 20: Destination | M2M2_ADDR_SENSOR_ADPD4000 = 0xC110
|  | 20: Request Command | M2M2_SENSOR_ADPD_COMMAND_CLOCK_CAL_REQ = 0x44 | Calibrate the 32M and 32K clock to reduce the deviation to a minimum
| PC to Device | 20: Response Command | M2M2_SENSOR_ADPD_COMMAND_CLOCK_CAL_RESP = 0x45
|  | 20: Status | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) | Clock calibration done successfully
| Header | [13 empty] | M2M2_APP_COMMON_STATUS_ERROR = 0x1 (or) | Error in clock claibration
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  |  |  | Device ID |  | 6-> in DVT1 Watches; 2 -> DVT2 Watches
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 00 | 08 | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11
| Description | 1: Command; 2: Status; 3: Device ID
| Data | 1: 44; 2: 00; 3: 06
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 10 | C7 | 05 | 00 | 08 | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11
| Description | 1: Command; 2: Status; 3: Device ID
| Data | 1: 45; 2: 00; 3: 06
Set PPG Library Configuration | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_MED_PPG = 0xC300
[13 empty] | Request Command |  | M2M2_APP_COMMON_CMD_SET_LCFG_REQ = 0x14 | Command to Set the PPG LCFG configuration
| PC to Device | [11 empty] | Response Command |  | M2M2_APP_COMMON_CMD_SET_LCFG_RESP = 0x15
[13 empty] | PPG LCFG ID |  | M2M2_SENSOR_PPG_LCFG_ID_ADPD4000 = 0x28
| Header | [11 empty] | Status |  | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) | PPG LCFG set successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | [5 empty] | M2M2_APP_COMMON_STATUS_ERROR = 1x0 (or) | Error in setting PPG LCFG
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 00 | 00 | 08 | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11
| Description | 1: Command; 2: Status; 3: PPG LCFG ID
| Data | 1: 14; 2: 00; 3: 28
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 00 | C7 | 05 | 00 | 0A | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10
| Description | 1: Command; 2: Status
| Data | 1: 15; 2: 00
Start PPG Application | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_MED_PPG = 0xC300
[13 empty] | Request Command |  | M2M2_APP_COMMON_CMD_STREAM_START_REQ = 0x4 | Command to start sensor
| PC to Device | [11 empty] | Response Command |  | M2M2_APP_COMMON_CMD_STREAM_START_RESP = 0x5
[13 empty] | Status |  | M2M2_APP_COMMON_STATUS_STREAM_STARTED = 0x2 | Stream started successfully
| Header | [13 empty] | M2M2_APP_COMMON_STATUS_STREAM_IN_PROGRESS = 0x4 | Already Stream started
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | [5 empty] | M2M2_APP_COMMON_STATUS_STREAM_NOT_STARTED = 0x7 | Error in starting stream
| Description | Source LSB | Source MSB | Destination LSB | Destination MSB | Length LSB | Length MSB | Checksum LSB | Checksum MSB
| Data | C7 | 05 | C3 | 00 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 04; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source LSB | Source MSB | Destination LSB | Destination MSB | Length LSB | Length MSB | Checksum LSB | Checksum MSB
| Data | C3 | 00 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 05; 2: 02; 3: 00; 4: 00
Subscribe PPG Stream | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_MED_PPG= 0xC300
[13 empty] | Request Command |  | M2M2_APP_COMMON_CMD_STREAM_SUBSCRIBE_REQ = 0xC | Command to subscribe the stream
| PC to Device | [11 empty] | Response Command |  | M2M2_APP_COMMON_CMD_STREAM_SUBSCRIBE_RESP = 0xD
[13 empty] | Stream address |  | M2M2_ADDR_MED_PPG_STREAM = 0xC400
| Header | [11 empty] | Status |  | M2M2_APP_COMMON_STATUS_SUBSCRIBER_ADDED = 0x9 | Subscriber added successfully to the stream mentioned
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 00 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0C; 2: 00; 3: 00; 4: C4
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 00 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0D; 2: 09; 3: 00; 4: C4

## PpgStop
PPG Stop Commands
UnSubscribe PPG Stream | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_MED_PPG= 0xC300
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_UNSUBSCRIBE_REQ = 0xE | Command to unsubscribe the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_UNSUBSCRIBE_RESP = 0xF
[10 empty] | Stream address | M2M2_ADDR_MED_PPG_STREAM = 0xC400
| Header | [8 empty] | Status | M2M2_APP_COMMON_STATUS_SUBSCRIBER_REMOVED = 0xA | Stream unsubscribed successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_SUBSCRIBER_COUNT_DECREMENT = 0xB | If stream subscribed more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 00 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0E; 2: 00; 3: 00; 4: C4
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 00 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0F; 2: 0A; 3: 00; 4: C4
Stop PPG Application | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_MED_PPG = 0xC300
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_STOP_REQ = 0x6 | Command to stop the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_STOP_RESP = 0x7
[10 empty] | Status | M2M2_APP_COMMON_STATUS_ERROR = 0x1
| Header | [9 empty] | M2M2_APP_COMMON_STATUS_STREAM_STOPPED = 0x3 | Stream stopped successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_STREAM_COUNT_DECREMENT =0x 6 | If stream started more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 00 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 06; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 00 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 07; 2: 03; 3: 00; 4: 00

## PpgData
Sample PPG data from CLI
RX::c4:00:c7:05:00:2e:00:00:28:43:15:00:3e:9d:32:00:07:00:53:05:b0:00:01:00:00:00:00:00:6e:00:00:00:6f:00:6d:92:3f:30:05:63:0f:10:c8:00:34:1c
PPG Data Stream | [9 empty] | source |  | M2M2_ADDR_MED_PPG_STREAM = 0xC400
[10 empty] | Destination |  | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Command |  | M2M2_SENSOR_COMMON_CMD_STREAM_DATA = 0x28
Device to PC
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C4 | 00 | C7 | 05 | 00 | 2E | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
Description | Command | Status | Sequence Number LSB | Sequence Number MSB | Timestamp LSB | Timestamp | Timestamp | Timestamp MSB | ADPD LibState LSB | ADPD LibState MSB | HR LSB | HR MSB
Data | 28 | 43 | 15 | 00 | 3E | 9D | 32 | 00 | 07 | 00 | 53 | 05
Payload
Byte Number | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32
Description | Confidence LSB | Confidence MSB | HR Type LSB | HR Type MSB | RR Interval LSB | RR interval MSB | Debug Info[0] LSB | Debug Info[0] MSB | Debug Info[1] LSB | Debug Info[1] MSB | Debug Info[2] LSB | Debug Info[2] MSB
Data | B0 | 00 | 01 | 00 | 00 | 00 | 00 | 00 | 6E | 00 | 00 | 00
Payload
Byte Number | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44
Description | Debug Info[3] LSB | Debug Info[3] MSB | Debug Info[4] LSB | Debug Info[4] MSB | Debug Info[5] LSB | Debug Info[5] MSB | Debug Info[6] LSB | Debug Info[6] MSB | Debug Info[7] LSB | Debug Info[7] MSB | Debug Info[8] LSB | Debug Info[8] MSB
Data | 6F | 00 | 6D | 92 | 3F | 30 | 05 | 63 | 0F | 10 | C8 | 00
Payload
Byte Number | 45 | 46
Description | Debug Info[9] LSB | Debug Info[9] MSB
Data | 34 | 1C

## PPG_CFG RdWt
PPG Configuration Read and Write Command
PC to Device
|  | Read PPG CFG 0x3 | [6 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
Header | [8 empty] | Destination | M2M2_ADDR_MED_PPG = 0xC300
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | Request Command | M2M2_APP_COMMON_CMD_READ_LCFG_REQ = 0x16 | Command to Read the PPG LCFG configuration
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB | Response Command | M2M2_APP_COMMON_CMD_READ_LCFG_RESP = 0x17
Data | C7 | 05 | C3 | 00 | 00 | 10 | 00 | 00 | PPG Configure Address | address = 3
Payload | [8 empty] | Status | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) | PPG LCFG Read successfully
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 |  | M2M2_APP_COMMON_STATUS_ERROR = 1x0 (or) | Error in Reading PPG LCFG
Description | Command | Status | Num_ops | Address | Value LSB | Value | Value | Value MSB
Data | 16 | 00 | 01 | 03 | 00 | 00 | 00 | 00
Device to PC
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C3 | 00 | C7 | 05 | 00 | 10 | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16
Description | Command | Status | Num_ops | Address | Value LSB | Value | Value | Value MSB
Data | 17 | 00 | 01 | 03 | 77 | 0a | 00 | 00
PC to Device
|  | Write PPG CFG 0x3 = 0xEF7 | [6 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
Header | [8 empty] | Destination | M2M2_ADDR_MED_PPG = 0xC300
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | Request Command | M2M2_APP_COMMON_CMD_WRITE_LCFG_REQ = 0x18 | Command to Read the PPG LCFG configuration
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB | Response Command | M2M2_APP_COMMON_CMD_WRITE_LCFG_RESP = 0x19
Data | C7 | 05 | C3 | 00 | 00 | 10 | 00 | 00 | Address, value | address = 3, value = 0x0EF7
Payload | [8 empty] | Status | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) | PPG LCFG Read successfully
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 |  | M2M2_APP_COMMON_STATUS_ERROR = 1x0 (or) | Error in Reading PPG LCFG
Description | Command | Status | Num_ops | Address | Value LSB | Value | Value | Value MSB
Data | 18 | 00 | 01 | 03 | F7 | 0E | 00 | 00
Device to PC
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C3 | 00 | C7 | 05 | 00 | 10 | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16
Description | Command | Status | Num_ops | Address | Value LSB | Value | Value | Value MSB
Data | 19 | 00 | 01 | 03 | F7 | 0E | 00 | 00

## SyncData
Sample SYNC ADPD-ADXl data from CLI
RX::c4:05:c7:05:00:48:00:00:28:00:c6:00:e8:80:32:00:11:84:32:00:81:02:81:02:80:02:81:02:82:02:82:02:5c:00:00:00:61:00:00:00:63:00:00:00:6d:00:00:00:10:00:0c:00:1c:00:18:00:27:00:23:00:27:00:27:00:43:01:44:01:43:01:43:01
SYNC ADPD-ADXL Data Stream | [9 empty] | source |  | M2M2_ADDR_MED_SYNC_ADPD_ADXL_STREAM = 0xC405
[10 empty] | Destination |  | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Command |  | M2M2_SENSOR_COMMON_CMD_STREAM_DATA = 0x28
Device to PC
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C4 | 05 | C7 | 05 | 00 | 48 | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
Description | Command | Status | Sequence Number LSB | Sequence Number MSB | PPG Timestamp LSB | PPG Timestamp | PPG Timestamp | PPG Timestamp MSB | ADXL Timestamp LSB | ADXL Timestamp | ADXL Timestamp | ADXL Timestamp MSB
Data | 28 | 00 | C6 | 00 | E8 | 80 | 32 | 00 | 11 | 84 | 32 | 00
Payload
Byte Number | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32
Description | Inc PPG Timestamp LSB | Inc PPG Timestamp MSB | Inc PPG Timestamp LSB | Inc PPG Timestamp MSB | Inc PPG Timestamp LSB | Inc PPG Timestamp MSB | Inc ADXL Timestamp LSB | Inc ADXL Timestamp MSB | Inc ADXl Timestamp LSB | Inc ADXL Timestamp MSB | Inc ADXL Timestamp LSB | Inc ADXL Timestamp MSB
Data | 81 | 02 | 81 | 02 | 80 | 02 | 81 | 02 | 82 | 02 | 82 | 02
Payload
Byte Number | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44
Description | PPG DATA LSB | PPG DATA | PPG DATA | PPG DATA MSB | PPG DATA LSB | PPG DATA | PPG DATA | PPG DATA MSB | PPG DATA LSB | PPG DATA | PPG DATA | PPG DATA MSB
Data | 5C | 00 | 00 | 00 | 61 | 00 | 00 | 00 | 63 | 00 | 00 | 00
Payload
Byte Number | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56
Description | PPG DATA LSB | PPG DATA | PPG DATA | PPG DATA MSB | ADXL - X DATA LSB | ADXL - X DATA MSB | ADXL - X DATA LSB | ADXL - X DATA MSB | ADXL - X DATA LSB | ADXL - X DATA MSB | ADXL - X DATA LSB | ADXL - X DATA MSB
Data | 6D | 00 | 00 | 00 | 10 | 00 | 0C | 00 | 1C | 00 | 18 | 00
Payload
Byte Number | 57 | 58 | 59 | 60 | 61 | 62 | 63 | 64 | 65 | 66 | 67 | 68
Description | ADXL - Y DATA LSB | ADXL - Y DATA MSB | ADXL - Y DATA LSB | ADXL - Y DATA MSB | ADXL - Y DATA LSB | ADXL - Y DATA MSB | ADXL - Y DATA LSB | ADXL - Y DATA MSB | ADXL - Z DATA LSB | ADXL - Z DATA MSB | ADXL - Z DATA LSB | ADXL - Z DATA MSB
Data | 27 | 00 | 23 | 00 | 27 | 00 | 27 | 00 | 43 | 01 | 44 | 01
Payload
Byte Number | 69 | 70 | 71 | 72
Description | ADXL - Z DATA LSB | ADXL - Z DATA MSB | ADXL - Z DATA LSB | ADXL - Z DATA MSB
Data | 43 | 01 | 43 | 01

## EdaStart
EDA Start Commands
Start EDA Sensor | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_MED_EDA = 0xC302
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_START_REQ = 0x4 | Command to start sensor
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_START_RESP = 0x5
[10 empty] | Status | M2M2_APP_COMMON_STATUS_STREAM_STARTED = 0x2 | Stream started successfully
| Header | [9 empty] | M2M2_APP_COMMON_STATUS_STREAM_IN_PROGRESS = 0x4 | Already Stream started
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_STREAM_NOT_STARTED = 0x7 | Error in starting stream
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 02 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 04; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 02 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 05; 2: 02; 3: 00; 4: 00
Subscribe EDA Stream | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_MED_EDA = 0xC302
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_SUBSCRIBE_REQ = 0xC | Command to subscribe the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_SUBSCRIBE_RESP = 0xD
[10 empty] | Stream address | M2M2_ADDR_MED_EDA_STREAM = 0xC402
| Header | [8 empty] | Status | M2M2_APP_COMMON_STATUS_SUBSCRIBER_ADDED = 0x9 | Subscriber added successfully to the stream mentioned
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 02 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0C; 2: 00; 3: 02; 4: C4
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source LSB | Source MSB | Destination LSB | Destination MSB | Length LSB | Length MSB | Checksum LSB | Checksum MSB
| Data | C3 | 02 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0D; 2: 09; 3: 02; 4: C4

## EdaStop
EDA Stop Commands
UnSubscribe EDA Stream | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_MED_EDA = 0xC302
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_UNSUBSCRIBE_REQ = 0xE | Command to unsubscribe the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_UNSUBSCRIBE_RESP = 0xF
[10 empty] | Stream address | M2M2_ADDR_MED_EDA_STREAM = 0xC402
| Header | [8 empty] | Status | M2M2_APP_COMMON_STATUS_SUBSCRIBER_REMOVED = 0xA | Stream unsubscribed successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_SUBSCRIBER_COUNT_DECREMENT = 0xB | If stream subscribed more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 02 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0E; 2: 00; 3: 02; 4: C4
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 02 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0F; 2: 0A; 3: 02; 4: C4
Stop EDA Sensor | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_MED_EDA = 0xC302
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_STOP_REQ = 0x6 | Command to stop the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_STOP_RESP = 0x7
[10 empty] | Status | M2M2_APP_COMMON_STATUS_ERROR = 0x1
| Header | [9 empty] | M2M2_APP_COMMON_STATUS_STREAM_STOPPED = 0x3 | Stream stopped successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_STREAM_COUNT_DECREMENT =0x 6 | If stream started more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 02 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 06; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 02 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 07; 2: 03; 3: 00; 4: 00

## EdaData
Sample EDA data from CLI
RX::c4:02:c7:05:00:3d:00:00:28:42:07:00:00:1f:ed:0a:00:aa:bb:2d:e9:3a:fc:0a:00:ab:ba:56:e7:55:0b:0b:00:af:ba:e4:e7:70:1a:0b:00:8f:bb:4d:e6:8b:29:0b:00:c2:bc:1d:e5:a7:38:0b:00:b3:bc:bd:e3
EDA Data Stream | [8 empty] | source | M2M2_ADDR_MED_EDA_STREAM = 0xC402
[9 empty] | Destination | M2M2_ADDR_APP_CLI = 0xC705
[9 empty] | Command | M2M2_SENSOR_COMMON_CMD_STREAM_DATA = 0x28
Device to PC
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C4 | 02 | C7 | 05 | 00 | 3D | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13
Description | Command | Status | Sequence Number LSB | Sequence Number MSB | DATA Type
Data | 28 | 42 | 07 | 00 | 00
Payload
Byte Number | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
Description | Timestamp LSB | Timestamp | Timestamp | Timestamp MSB | Real DATA LSB | Real DATA MSB | Imaginary DATA LSB | Imaginary DATA MSB
Data | 1F | ED | 0A | 00 | AA | BB | 2D | E9
Payload
Byte Number | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29
Description | Timestamp LSB | Timestamp | Timestamp | Timestamp MSB | Real DATA LSB | Real DATA MSB | Imaginary DATA LSB | Imaginary DATA MSB
Data | 3A | FC | 0A | 00 | AB | BA | 56 | E7
Payload
Byte Number | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37
Description | Timestamp LSB | Timestamp | Timestamp | Timestamp MSB | Real DATA LSB | Real DATA MSB | Imaginary DATA LSB | Imaginary DATA MSB
Data | 55 | 0B | 0B | 00 | AF | BA | E4 | E7
Payload
Byte Number | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45
Description | Timestamp LSB | Timestamp | Timestamp | Timestamp MSB | Real DATA LSB | Real DATA MSB | Imaginary DATA LSB | Imaginary DATA MSB
Data | 70 | 1A | 0B | 00 | 8F | BB | 4D | E6
Payload
Byte Number | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53
Description | Timestamp LSB | Timestamp | Timestamp | Timestamp MSB | Real DATA LSB | Real DATA MSB | Imaginary DATA LSB | Imaginary DATA MSB
Data | 8B | 29 | 0B | 00 | C2 | BC | 1D | E5
Payload
Byte Number | 54 | 55 | 56 | 57 | 58 | 59 | 60 | 61
Description | Timestamp LSB | Timestamp | Timestamp | Timestamp MSB | Real DATA LSB | Real DATA MSB | Imaginary DATA LSB | Imaginary DATA MSB
Data | A7 | 38 | 0B | 00 | B3 | BC | BD | E3

## AdpdStart
ADPD Start Commands
[14 empty] | 4 | Give Turn AGC On or Off Command; command to be given for the slots/LEDs separately
|  |  |  | 1 | Load Adpd Configuration with Device ID as required by the Slot or LED required to be operated | [8 empty] | 5 | Give Set Slot command separately, to configure the Slots for data format & channel selection as required.(Step 5 maybe skipped, since Load Adpd Configuration command confires the device as required)
|  |  |  | 2 | Clock Calibration with clock ID based on DVT 1/ 2 of the Watch | [8 empty] | 6 | Start the ADPD Sensor
|  |  |  | 3 | Change Freq of ADPD, if required, with reg write command | [8 empty] | 7 | Give Subscribe command to all the slots separately, to start receiving data from it
Load Adpd configuration | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_SENSOR_ADPD4000 = 0xC110
[13 empty] | Request Command |  | M2M2_SENSOR_ADPD_COMMAND_LOAD_CFG_REQ = 0x42 |  | Command to Load Adpd device configuration
| PC to Device | [11 empty] | Response Command |  | M2M2_SENSOR_ADPD_COMMAND_LOAD_CFG_RESP = 0x43
[13 empty] | Device ID |  | M2M2_SENSOR_ADPD4000_DEVICE_4000_G = 0x28 |  | For Green LED:M2M2_SENSOR_ADPD4000_DEVICE_4000_G = 0x28 For Red LED:M2M2_SENSOR_ADPD4000_DEVICE_4000_R = 0x29 For IR LED:M2M2_SENSOR_ADPD4000_DEVICE_4000_IR = 0x2A For Blue LED:M2M2_SENSOR_ADPD4000_DEVICE_4000_B = 0x2B For G, R, IR, B LEDs: M2M2_SENSOR_ADPD4000_DEVICE_4000_G_R_IR_B = 0x2C
| Header | [11 empty] | Status |  | M2M2_APP_COMMON_STATUS_OK = 0 (or) |  | ADPD device loaded successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | [5 empty] | M2M2_APP_COMMON_STATUS_ERROR = 1 (or) |  | Error in Loading ADPD device
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 00 | 15 | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
| Description | Command | Status | Device ID | return data | return data | return data | return data | return data | return data | return data | return data | return data | return data
| Data | 42 | 00 | 28 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 10 | C7 | 05 | 00 | 15 | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
| Description | Command | Status | Device ID | return data | return data | return data | return data | return data | return data | return data | return data | return data | return data
| Data | 43 | 00 | 28 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00
Clock Calibration |  | 20: source | M2M2_ADDR_APP_CLI = 0xC705
|  | 20: Destination | M2M2_ADDR_SENSOR_ADPD4000 = 0xC110
|  | 20: Request Command | M2M2_SENSOR_ADPD_COMMAND_CLOCK_CAL_REQ = 0x44 |  | Calibrate the 32M and 32K clock to reduce the deviation to a minimum
| PC to Device | 20: Response Command | M2M2_SENSOR_ADPD_COMMAND_CLOCK_CAL_RESP = 0x45
|  | 20: Status | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) |  | Clock calibration done successfully
| Header | [13 empty] | M2M2_APP_COMMON_STATUS_ERROR = 0x1 (or) |  | Error in clock claibration
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  |  |  | clock cal id |  | 6-> in DVT1 Watches; 2 -> DVT2 Watches
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 00 | 0B | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11
| Description | 1: Command; 2: Status; 3: clockcalid
| Data | 1: 44; 2: 00; 3: 02
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 10 | C7 | 05 | 00 | B | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11
| Description | 1: Command; 2: Status; 3: clockcalid
| Data | 1: 44; 2: 00; 3: 02
Register Write Command
PC to Device
| Write ADPD register 0xD to update the Frequency of ADPD: 0x4E20 -> 50Hz; 0x2710 -> 100Hz; 07D0 -> 500Hz | [8 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
Header | [9 empty] | Destination |  | M2M2_ADDR_SENSOR_ADPD4000 = 0xC110
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | Request Command |  | M2M2_SENSOR_COMMON_CMD_WRITE_REG_16_REQ = 0x23 |  |  |  | Command to Perform a register write operation on a device.
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB |  | Response Command |  | M2M2_SENSOR_COMMON_CMD_WRITE_REG_16_RESP = 0x24
Data | C7 | 05 | C1 | 10 | 00 | 0F | 00 | 00 |  | Status |  | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) |  |  |  | Register read/write operation performed successfully
Payload | [11 empty] | M2M2_APP_COMMON_STATUS_ERROR = 0x1 (or) |  |  |  | Error in Read/Write register opertaion
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 14
Description | Command | Status | Num_ops | Address LSB | Address MSB | Value LSB | Value MSB
Data | 23 | 00 | 01 | 0D | 00 | 10 | 27
Device to PC
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C1 | 10 | C7 | 05 | 00 | 0F | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 14
Description | Command | Status | Num_ops | Address LSB | Address MSB | Value LSB | Value MSB
Data | 24 | 00 | 01 | 0D | 00 | 10 | 27
Set ADPD Slot | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_SENSOR_ADPD4000_DEVICE_4000_G = 0xC110
[13 empty] | Request Command |  | M2M2_SENSOR_ADPD_COMMAND_SET_SLOT_REQ = 0x46 |  | Command to set Adpd device slot
[13 empty] | ADPD Slot Number |  | 0x6 -> Slot F with Green LED |  | Slots A to L are given with values from 0x1 to 0xC 0x6 -> Slot F with Green LED 0x7 -> Slot G with Red LED 0x8 -> Slot H with IR LED 0x9 -> Slot I with Blue LED
[13 empty] | ADPD Slot Enable |  | 1 |  | 0x0 -> To disable the Slot 0x1 -> To enable the Slot
[13 empty] | Slot Format |  | 0x4 |  | 4 byte format
[13 empty] | Channel Number |  | 0x3 |  | 0x1 -> One channel enable 0x3 -> Two channel enable
| PC to Device | [11 empty] | Response Command |  | M2M2_SENSOR_ADPD_COMMAND_SET_SLOT_RESP = 0x47
[13 empty] | Status |  | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) |  | ADPD slot set successfully
| Header | [13 empty] | M2M2_APP_COMMON_STATUS_ERROR = 0x1 (or) |  | Error in setting ADPD slot
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 00 | 0F | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15
| Description | Command | Status | ADPD Slot Num | ADPD Slot Enable | Slot Format LSB | Slot Format MSB | Channel Number
| Data | 46 | 00 | 06 | 1 | 04 | 00 | 03
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 10 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15
| Description | Command | Status | ADPD Slot Num | ADPD Slot Enable | Slot Format LSB | Slot Format MSB | Channel Number
| Data | 47 | 00 | 06 | 1 | 04 | 00 | 03
TurnAGC ON or OFF | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_SENSOR_ADPD4000_DEVICE_4000_G = 0xC110
[13 empty] | Request Command |  | M2M2_SENSOR_ADPD_COMMAND_AGC_ON_OFF_REQ = 0x76 |  | Command to set Adpd device slot
[13 empty] | agc_type |  | 0x1 |  | To Select between green/blue/red/ir/mwl(all 4 LEDs) mwl = 0 g = 1 r = 2 ir = 3 b = 4
[13 empty] | agc_cntrl |  | 0x1 |  | To Enable AGC, agc_cntrl= 1 To Disable AGC, agc_cntrl= = 0
[13 empty] | Response Command |  | M2M2_SENSOR_ADPD_COMMAND_AGC_ON_OFF_REQ = 0x77
[13 empty] | Status |  | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) |  | AGC Control set successfully
[15 empty] | M2M2_APP_COMMON_STATUS_ERROR = 0x1 (or) |  | Error in setting AGC Control
PC to Device
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 00 | 0D | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 12
| Description | Command | Status | num_of_ops | agc_cntrl | agc_type
| Data | 76 | 00 | 01 | 01 | 01
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 01 | C7 | 05 | 00 | 0D | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 12
| Description | Command | Status | num_of_ops | agc_cntrl | agc_type
| Data | 77 | 00 | 01 | 01 | 01
Start ADPD sensor | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_SENSOR_ADPD4000 = 0xC110
[13 empty] | Request Command |  | M2M2_APP_COMMON_CMD_STREAM_START_REQ = 0x4 |  | Command to start sensor
| PC to Device | [11 empty] | Response Command |  | M2M2_APP_COMMON_CMD_STREAM_START_RESP = 0x5
[13 empty] | Status |  | M2M2_APP_COMMON_STATUS_STREAM_STARTED = 0x2 |  | Stream started successfully
| Header | [13 empty] | M2M2_APP_COMMON_STATUS_STREAM_IN_PROGRESS = 0x4 |  | Already Stream started
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | [5 empty] | M2M2_APP_COMMON_STATUS_STREAM_NOT_STARTED = 0x7 |  | Error in starting stream
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 04; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 10 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 05; 2: 02; 3: 00; 4: 00
[13 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
Subscribe ADPD Stream | [12 empty] | Destination |  | M2M2_ADDR_SENSOR_ADPD4000 = 0xC110
[13 empty] | Request Command |  | M2M2_APP_COMMON_CMD_STREAM_SUBSCRIBE_REQ = 0xC
[13 empty] | Response Command |  | M2M2_APP_COMMON_CMD_STREAM_SUBSCRIBE_RESP = 0xD |  | Command to subscribe the stream
| PC to Device | [11 empty] | Stream address |  | M2M2_ADDR_SENSOR_ADPD_STREAM6 = 0xC216 --> Green LED M2M2_ADDR_SENSOR_ADPD_STREAM7 = 0xC217 --> Red LED M2M2_ADDR_SENSOR_ADPD_STREAM8 = 0xC218 --> IR LED M2M2_ADDR_SENSOR_ADPD_STREAM9 = 0xC219 --> Blue LED
[13 empty] | Status |  | M2M2_APP_COMMON_STATUS_SUBSCRIBER_ADDED = 0x9
| Header | [15 empty] | Subscriber added successfully to the stream mentioned
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0C; 2: 00; 3: 16; 4: C2
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 10 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0D; 2: 09; 3: 16; 4: C2

## AdpdStop
ADPD Stop Commands
[6 empty] | 1 | Give UnSubscribe command to all the slots separately, to stop receiving data from it. All slots subscribed needs to be unsubscribed
[6 empty] | 2 | Stop the ADPD Sensor
UnSubscribe ADPD Stream | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_SENSOR_ADPD4000 = 0xC110
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_UNSUBSCRIBE_REQ = 0xE | Command to unsubscribe the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_UNSUBSCRIBE_RESP = 0xF
[10 empty] | Stream address | M2M2_ADDR_SENSOR_ADPD_STREAM6 = 0xC216
| Header | [8 empty] | Status | M2M2_APP_COMMON_STATUS_SUBSCRIBER_REMOVED = 0xA | Stream unsubscribed successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_SUBSCRIBER_COUNT_DECREMENT = 0xB | If stream subscribed more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0E; 2: 00; 3: 16; 4: C2
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 10 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0F; 2: 0A; 3: 16; 4: C2
Stop ADPD Sensor | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_SENSOR_ADPD4000 = 0xC110
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_STOP_REQ = 0x6 | Command to stop the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_STOP_RESP = 0x7
[10 empty] | Status | M2M2_APP_COMMON_STATUS_ERROR = 0x1
| Header | [9 empty] | M2M2_APP_COMMON_STATUS_STREAM_STOPPED = 0x3 | Stream stopped successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_STREAM_COUNT_DECREMENT =0x6 | If stream started more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 06; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 10 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 07; 2: 03; 3: 00; 4: 00

## AdpdData
Sample ADPD data from CLI
RX::c2:16:c8:08:00:2c:00:00:28:00:8d:00:04:00:02:8e:6f:2f:76:02:61:c4:01:00:c4:c4:01:00:56:1f:00:00:58:1f:00:00:5b:1f:00:00:8c:1f:00:00
Frequently Asked Questions
1. What is the meaning of Channel corresponding to each wavelength and/or sampling rate. Do you have any document explaining these? Ans: The Green LED data is coming in two channels of Slot F of ADPD4100. Slot Assignment for ADPD on Study Watch: ----------------------------------------------------- A - ECG from ADPD D - Thermistor Temperature E - Calibration Resistor F - Green LED G - Red LED H - IR LED I - Blue LED Please check the ADPD4x00 datasheet shared, where its describes about the 12 slots of ADPD sensor with 2 channels each.
2. Can you please explain the differences between Ch01 and Ch02? Which one should we take if we need IR ppg sample. Ans: There is no particular guidance on Ch1 or Ch2. You could use one of them based on its quality or may be take an average of them for further processing
3. What is the return value for the 0x0D response command for "Subscribe ADPD Stream" for error conditions? (as per the documentation success is 0x09) Ans: Currently there is no error code for this command.
4. What is the stream address for subscription of all streams for "Subscribe ADPD Stream"? If we want to subscribe to all streams do we need to call this multiple times with different stream addresses? Ans: Subscribe command needs to be given separately, to the streams required, with different stream address.
5. If the stream is subscribed more than one time (UnSubscribe ADPD Stream ( M2M2_APP_COMMON_STATUS_SUBSCRIBER_COUNT_DECREMENT = 0xB ) do we need to call Unsubscribe multiple times till we get 0xA? Ans: Yes. Basically in the stop sequence, unsubscribe all the streams which have been subscribed, one after the other.
6. The timestamp of all of these data are the same (millisec since epoch .. though we had set the time of the watch correctly from ntp). Ans: Timestamp in the packets is in ticks resolution from the firmware. You would need to convert this to millisec : 1 ms = 32768 ticks
7. BLEProviderADIStudyWatch.responseData[176] c216c700002c0000280034000400010d4fa88102d8a40000c8a3000000000000000000000000000000000000c216c700002c0000280034000400020d4fa88102b6bf0000aebe000000000000000000000000000000000000c216c700002c0000280035000400011554a8810236a3000046a3000000000000000000000000000000000000c216c700002c0000280035000400021554a8810211be000030be000000000000000000000000000000000000 When we subscribe, we receive a packet of length 176 which contains 4 AdpdData. First 2 AdpdData packets' seq_id are the same and next 2 AdpdData packets' seq id are the same. Can you explain? How do we know which sensor (IR, R ...) is sending these packets? Ans: You need to parse the above 4 packets received as mentioned in ADPDData sheet, with which you would get a Ch1, Ch2 data with same sequence number, then a Ch1, Ch2 data with next sequence number, from stream ID which belongs to Green LED slot. You would have got this since M2M2_ADDR_SENSOR_ADPD_STREAM6 was subscribed. Similarly you need to subscribe to IR, R .. to get data from these slots.
8. Please explain "First ADPD Data" of length 24. If this contains the little endian PPG (eg: Green) the signal does not look like a PPG as we know (refer to the plot which we emailed before). The plotted signal is collected when we wear the watch in the wrist or we touch the sensor with a finger tip. Ans: At 100 Hz, the array First ADPD Data[24] contains 6 ADPD samples of 4 bytes each. This needs to be parsed and plotted as 6 samples of that Channel data.
9. Can you explain the "Stop ADPD Sensor" use-case for error condition M2M2_APP_COMMON_STATUS_STREAM_COUNT_DECREMENT =0x6 ? As per the documentation, we can not start Adpd Sensor multiple times; if we send the start command multiple times, we will get M2M2_APP_COMMON_STATUS_STREAM_IN_PROGRESS = 0x4 error. Ans: While using other applications, ADPD sensor may be started, as internal to that application (eg: temperature app )in which case the Start ADPD Sensor status is expected to be already incremented. Thereafter, if you were to give this “Stop ADPD Sensor” you would see the M2M2_APP_COMMON_STATUS_STREAM_COUNT_DECREMENT Condition.
10. How is ADPD Data filled in First ADPD Data[24] array? Ans: Considering no decimation, it varies with the output data rate(ODR) configured if odr < 25Hz, no. of samples = 1 if odr < 75Hz, no. of samples = 2 else, no. of samples = 6
| ADPD Data Stream | [8 empty] | source | M2M2_ADDR_SENSOR_ADPD_STREAM6 = 0xC216
[10 empty] | Destination | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Command | M2M2_SENSOR_COMMON_CMD_STREAM_DATA = 0x28
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C2 | 16 | C7 | 05 | 00 | 2C | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15
| Description | Command | Status | Sequence Number LSB | Sequence Number MSB | DATA Type LSB | DATA Type MSB | Channel Number
| Data | 28 | 00 | 8D | 00 | 04 | 00 | 02
Payload
| Byte Number | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23
| Description | Timestamp LSB | Timestamp | Timestamp | Timestamp MSB | sample_num | First ADPD Data LSB | First ADPD Data | First ADPD Data
| Data | 8e | 6f | 2f | 76 | 02 | 61 | c4 | 01
Payload
| Byte Number | 24 | 25 | 26 | 27 | 28 | 29
| Description | First ADPD Data | First ADPD Data | First ADPD Data | First ADPD Data | First ADPD Data | First ADPD Data
| Data | 00 | c4 | c4 | 01 | 00 | 56
Payload
| Byte Number | 30 | 31 | 32 | 33 | 34 | 35
| Description | First ADPD Data | First ADPD Data | First ADPD Data | First ADPD Data | First ADPD Data | First ADPD Data
| Data | 1f | 00 | 00 | 58 | 1f | 00
Payload
| Byte Number | 36 | 37 | 38 | 39 | 40 | 41
| Description | First ADPD Data | First ADPD Data | First ADPD Data | First ADPD Data | First ADPD Data | First ADPD Data
| Data | 00 | 5b | 1f | 00 | 00 | 8c
Payload
| Byte Number | 16: 42; 17: 43; 18: 44
| Description | 16: First ADPD Data; 17: First ADPD Data; 18: First ADPD Data MSB
| Data | 16: 1f; 17: 00; 18: 00

## ADPD_logging_start
ADPD Logging Start Commands
|  |  |  | 1 | Subscribe for ADPD streams from FS | [8 empty] | 5 | Give Turn AGC On or Off Command; command to be given for the slots/LEDs separately
|  |  |  | 2 | Load Adpd Configuration with Device ID as required by the Slot or LED required to be operated | [8 empty] | 6 | Start the ADPD Sensor
|  |  |  | 3 | Clock Calibration with clock ID based on DVT 1/ 2 of the Watch | [8 empty] | 7 | Start FS logging
|  |  |  | 4 | Change Freq of ADPD, if required, with reg write command
FS Subscribe ADPD Stream | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_SYS_FS= 0xC501
[13 empty] | Request Command |  | M2M2_FILE_SYS_CMD_LOG_STREAM_REQ = 0x72 |  | Command to subscribe the stream
| PC to Device | [11 empty] | Response Command |  | M2M2_FILE_SYS_CMD_LOG_STREAM_RESP = 0x73
[13 empty] | Stream address |  | M2M2_ADDR_SENSOR_ADPD_STREAM6 = 0xC216 --> Green LED M2M2_ADDR_SENSOR_ADPD_STREAM7 = 0xC217 --> Red LED M2M2_ADDR_SENSOR_ADPD_STREAM8 = 0xC218 --> IR LED M2M2_ADDR_SENSOR_ADPD_STREAM9 = 0xC219 --> Blue LED
| Header | [11 empty] | Status |  | M2M2_APP_COMMON_STATUS_SUBSCRIBER_ADDED = 0x9 |  | Subscriber added successfully to the stream mentioned
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C5 | 01 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 72; 2: 00; 3: 16; 4: C2
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 01 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 73; 2: 09; 3: 16; 4: C2
Load Adpd configuration | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_SENSOR_ADPD4000 = 0xC110
[13 empty] | Request Command |  | M2M2_SENSOR_ADPD_COMMAND_LOAD_CFG_REQ = 0x42 |  | Command to Load Adpd device configuration
| PC to Device | [11 empty] | Response Command |  | M2M2_SENSOR_ADPD_COMMAND_LOAD_CFG_RESP = 0x43
[13 empty] | Device ID |  | M2M2_SENSOR_ADPD4000_DEVICE_4000_G = 0x28 |  | For Green LED:M2M2_SENSOR_ADPD4000_DEVICE_4000_G = 0x28 For Red LED:M2M2_SENSOR_ADPD4000_DEVICE_4000_R = 0x29 For IR LED:M2M2_SENSOR_ADPD4000_DEVICE_4000_IR = 0x2A For Blue LED:M2M2_SENSOR_ADPD4000_DEVICE_4000_B = 0x2B For G, R, IR, B LEDs: M2M2_SENSOR_ADPD4000_DEVICE_4000_G_R_IR_B = 0x2C
| Header | [11 empty] | Status |  | M2M2_APP_COMMON_STATUS_OK = 0 (or) |  | ADPD device loaded successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | [5 empty] | M2M2_APP_COMMON_STATUS_ERROR = 1 (or) |  | Error in Loading ADPD device
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 00 | 15 | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
| Description | Command | Status | Device ID | return data | return data | return data | return data | return data | return data | return data | return data | return data | return data
| Data | 42 | 00 | 28 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 10 | C7 | 05 | 00 | 15 | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
| Description | Command | Status | Device ID | return data | return data | return data | return data | return data | return data | return data | return data | return data | return data
| Data | 43 | 00 | 28 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00
Clock Calibration |  | 20: source | M2M2_ADDR_APP_CLI = 0xC705
|  | 20: Destination | M2M2_ADDR_SENSOR_ADPD4000 = 0xC110
|  | 20: Request Command | M2M2_SENSOR_ADPD_COMMAND_CLOCK_CAL_REQ = 0x44 |  | Calibrate the 32M and 32K clock to reduce the deviation to a minimum
| PC to Device | 20: Response Command | M2M2_SENSOR_ADPD_COMMAND_CLOCK_CAL_RESP = 0x45
|  | 20: Status | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) |  | Clock calibration done successfully
| Header | [13 empty] | M2M2_APP_COMMON_STATUS_ERROR = 0x1 (or) |  | Error in clock claibration
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  |  |  | Device ID |  | 6-> in DVT1 Watches; 2 -> DVT2 Watches
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 00 | 0B | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11
| Description | 1: Command; 2: Status; 3: Device ID
| Data | 1: 44; 2: 00; 3: 02
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 10 | C7 | 05 | 00 | B | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11
| Description | 1: Command; 2: Status; 3: Device ID
| Data | 1: 44; 2: 00; 3: 02
Register Write Command
PC to Device
|  | 1: Write ADPD register 0xD to update the Frequency of ADPD: 0x4E20 -> 50Hz; 0x2710 -> 100Hz; 07D0 -> 500Hz |  | source |  | M2M2_ADDR_APP_CLI = 0xC705
| Header | [9 empty] | Destination |  | M2M2_ADDR_SENSOR_ADPD4000 = 0xC110
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | Request Command |  | M2M2_SENSOR_COMMON_CMD_WRITE_REG_16_REQ = 0x23 |  |  | Command to Perform a register write operation on a device.
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB |  | Response Command |  | M2M2_SENSOR_COMMON_CMD_WRITE_REG_16_RESP = 0x24
| Data | C7 | 05 | C1 | 10 | 00 | 0F | 00 | 00 |  | Status |  | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) |  |  | Register read/write operation performed successfully
| Payload | [11 empty] | M2M2_APP_COMMON_STATUS_ERROR = 0x1 (or) |  |  | Error in Read/Write register opertaion
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 14
| Description | Command | Status | Num_ops | Address LSB | Address MSB | Value LSB | Value MSB
| Data | 23 | 00 | 01 | 0D | 00 | 10 | 27
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 10 | C705 | C7 | 00 | 0F | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 14
| Description | Command | Status | Num_ops | Address LSB | Address MSB | Value LSB | Value MSB
| Data | 24 | 00 | 01 | 34 | 00 | 00 | 03
TurnAGC ON or OFF | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_SENSOR_ADPD4000_DEVICE_4000_G = 0xC110
[13 empty] | Request Command |  | M2M2_SENSOR_ADPD_COMMAND_AGC_ON_OFF_REQ = 0x76 |  | Command to set Adpd device slot
[13 empty] | agc_type |  | 0x1 |  | To Select between green/blue/red/ir/mwl(all 4 LEDs) mwl = 0 g = 1 r = 2 ir = 3 b = 4
[13 empty] | agc_cntrl |  | 0x1 |  | To Enable AGC, agc_cntrl= 1 To Disable AGC, agc_cntrl= = 0
[13 empty] | Response Command |  | M2M2_SENSOR_ADPD_COMMAND_AGC_ON_OFF_REQ = 0x77
[13 empty] | Status |  | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) |  | AGC Control set successfully
[15 empty] | M2M2_APP_COMMON_STATUS_ERROR = 0x1 (or) |  | Error in setting AGC Control
PC to Device
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 00 | 0D | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 12
| Description | Command | Status | num_of_ops | agc_cntrl | agc_type
| Data | 76 | 00 | 01 | 01 | 01
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 01 | C7 | 05 | 00 | 0D | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 12
| Description | Command | Status | num_of_ops | agc_cntrl | agc_type
| Data | 77 | 00 | 01 | 01 | 01
Start ADPD sensor | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_SENSOR_ADPD4000 = 0xC110
[13 empty] | Request Command |  | M2M2_APP_COMMON_CMD_STREAM_START_REQ = 0x4 |  | Command to start sensor
| PC to Device | [11 empty] | Response Command |  | M2M2_APP_COMMON_CMD_STREAM_START_RESP = 0x5
[13 empty] | Status |  | M2M2_APP_COMMON_STATUS_STREAM_STARTED = 0x2 |  | Stream started successfully
| Header | [13 empty] | M2M2_APP_COMMON_STATUS_STREAM_IN_PROGRESS = 0x4 |  | Already Stream started
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | [5 empty] | M2M2_APP_COMMON_STATUS_STREAM_NOT_STARTED = 0x7 |  | Error in starting stream
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 04; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 10 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 05; 2: 02; 3: 00; 4: 00
File open Command | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_SYS_FS= 0xC501
[13 empty] | Request Command |  | M2M2_FILE_SYS_CMD_START_LOGGING_REQ = 0x76 |  | Command to open new file
| PC to Device | [11 empty] | Response Command |  | M2M2_FILE_SYS_CMD_START_LOGGING_RESP = 0x77
[13 empty] | Status |  | M2M2_FILE_SYS_STATUS_OK = 0x41 (or) |  | File created successfully performed successfully
| Header | [13 empty] | M2M2_FILE_SYS_STATUS_LOGGING_IN_PROGRESS = 0x4E |  | Already File in open state, Close the existing file to create now one
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | [5 empty] | M2M2_FILE_SYS_STATUS_ERROR = 0x42 |  | Error in Creating file
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB | [5 empty] | M2M2_FILE_SYS_ERR_MAX_FILE_COUNT = 0x4A |  | Maximum file count reached(64), Format the file system to create new file.
| Data | C7 | 05 | C5 | 01 | 00 | 0C | 00 | 00 | [5 empty] | M2M2_FILE_SYS_ERR_INVALID = 0x45 |  | Invalid file directory
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 76; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 01 | C7 | 05 | 00 | 0A | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10
| Description | 1: Command; 2: Status
| Data | 1: 77; 2: 41

## ADPD_logging_stop
ADPD Logging Stop Commands
|  |  |  | 1 | Stop the ADPD Sensor
|  |  |  | 2 | UnSubscribe for all ADPD streams(subscribed earlier) from FS
|  |  |  | 3 | Stop FS logging
Stop ADPD Sensor | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_SENSOR_ADPD4000 = 0xC110
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_STOP_REQ = 0x6 | Command to stop the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_STOP_RESP = 0x7
[10 empty] | Status | M2M2_APP_COMMON_STATUS_ERROR = 0x1
| Header | [9 empty] | M2M2_APP_COMMON_STATUS_STREAM_STOPPED = 0x3 | Stream stopped successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_STREAM_COUNT_DECREMENT =0x 6 | If stream started more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 06; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 10 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 07; 2: 03; 3: 00; 4: 00
UnSubscribe FS ADPD Stream | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_SYS_FS= 0xC501
[10 empty] | Request Command | M2M2_FILE_SYS_CMD_STOP_STREAM_REQ = 0x74 | Command to unsubscribe the stream
| PC to Device | [8 empty] | Response Command | M2M2_FILE_SYS_CMD_STOP_STREAM_REQ = 0x75
[10 empty] | Stream address | M2M2_ADDR_SENSOR_ADPD_STREAM6 = 0xC216 --> Green LED M2M2_ADDR_SENSOR_ADPD_STREAM7 = 0xC217 --> Red LED M2M2_ADDR_SENSOR_ADPD_STREAM8 = 0xC218 --> IR LED M2M2_ADDR_SENSOR_ADPD_STREAM9 = 0xC219 --> Blue LED
| Header | [8 empty] | Status | M2M2_APP_COMMON_STATUS_SUBSCRIBER_REMOVED = 0xA | Stream unsubscribed successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_SUBSCRIBER_COUNT_DECREMENT = 0xB | If stream subscribed more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C5 | 01 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 74; 2: 00; 3: 16; 4: C2
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 01 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 75; 2: 0A; 3: 16; 4: C2
File Close Command | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_SYS_FS= 0xC501
[10 empty] | Request Command | M2M2_FILE_SYS_CMD_STOP_LOGGING_REQ = 0x78 | Command to close the file
| PC to Device | [8 empty] | Response Command | M2M2_FILE_SYS_CMD_STOP_LOGGING_RESP = 0x79
[10 empty] | Status | M2M2_FILE_SYS_STATUS_LOGGING_STOPPED = 0x4D (or) | File closed successfully performed successfully
| Header | [9 empty] | M2M2_FILE_SYS_STATUS_ERROR = 0x42 (or) | Error in Closing the file
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_FILE_SYS_STATUS_LOGGING_NOT_STARTED = 0x50 | Logging not yet started
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C5 | 01 | 00 | 0B | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11
| Description | 1: Command; 2: Status; 3: stop type
| Data | 1: 78; 2: 00; 3: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 01 | C7 | 05 | 00 | 0B | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11
| Description | 1: Command; 2: Status; 3: stop type
| Data | 1: 79; 2: 41; 3: 00

## AdxlStart
ADXL Start Commands
|  |  |  | 1 | start adxl
|  |  |  | 2 | Subscrib adxl
Start ADXL sensor | [11 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[12 empty] | Destination | M2M2_ADDR_SENSOR_ADXL = 0xC102
[12 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_START_REQ = 0x4 |  | Command to start sensor
| PC to Device | [10 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_START_RESP = 0x5
[12 empty] | Status | M2M2_APP_COMMON_STATUS_STREAM_STARTED = 0x2 |  | Stream started successfully
| Header | [11 empty] | M2M2_APP_COMMON_STATUS_STREAM_IN_PROGRESS = 0x4 |  | Already Stream started
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  |  |  | M2M2_APP_COMMON_STATUS_STREAM_NOT_STARTED = 0x7 |  | Error in starting stream
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 02 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 04; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 02 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 05; 2: 02; 3: 00; 4: 00
Register Write Command to set ADXL at 100Hz
PC to Device
|  | 1: Write ADXL register 0x2C | source | M2M2_ADDR_APP_CLI = 0xC705
Header | [9 empty] | Destination | M2M2_ADDR_SENSOR_ADXL = 0xC102
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | Request Command | M2M2_SENSOR_COMMON_CMD_WRITE_REG_16_REQ = 0x23 |  |  | Command to Perform a register operation on a device.
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB |  | Response Command | M2M2_SENSOR_COMMON_CMD_WRITE_REG_16_RESP = 0x24
Data | C7 | 05 | C1 | 2 | 00 | 0F | 00 | 00 |  | Status | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) |  |  | Register read/write operation performed successfully
Payload | [10 empty] | M2M2_APP_COMMON_STATUS_ERROR = 0x1 (or) |  |  | Error in Read/Write register opertaion
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 14
Description | Command | Status | Num_ops | Address LSB | Address MSB | Value LSB | Value MSB
Data | 23 | 00 | 01 | 2C | 00 | 9B | 00
Device to PC
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C1 | 2 | C7 | 05 | 00 | 0F | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 14
Description | Command | Status | Num_ops | Address LSB | Address MSB | Value LSB | Value MSB
Data | 24 | 00 | 01 | 2C | 00 | 9B | 00
Subscribe ADXL Stream | [11 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[12 empty] | Destination | M2M2_ADDR_SENSOR_ADXL = 0xC102
[12 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_SUBSCRIBE_REQ = 0xC |  | Command to subscribe the stream
| PC to Device | [10 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_SUBSCRIBE_RESP = 0xD
[12 empty] | Stream address | M2M2_ADDR_SENSOR_ADXL_STREAM = 0xC202
| Header | [10 empty] | Status | M2M2_APP_COMMON_STATUS_SUBSCRIBER_ADDED = 0x9 |  | Subscriber added successfully to the stream mentioned
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 02 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0C; 2: 00; 3: 02; 4: C2
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 02 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0D; 2: 09; 3: 02; 4: C2

## AdxlStop
ADXL Stop Commands
Stop ADXL Sensor | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_SENSOR_ADXL = 0xC102
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_STOP_REQ = 0x6 | Command to stop the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_STOP_RESP = 0x7
[10 empty] | Status | M2M2_APP_COMMON_STATUS_ERROR = 0x1
| Header | [9 empty] | M2M2_APP_COMMON_STATUS_STREAM_STOPPED = 0x3 | Stream stopped successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_STREAM_COUNT_DECREMENT =0x 6 | If stream started more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 02 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 06; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source LSB | Source MSB | Destination LSB | Destination MSB | Length LSB | Length MSB | Checksum LSB | Checksum MSB
| Data | C1 | 02 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 07; 2: 03; 3: 00; 4: 00
UnSubscribe ADXL Stream | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_SENSOR_ADXL = 0xC102
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_UNSUBSCRIBE_REQ = 0xE | Command to unsubscribe the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_UNSUBSCRIBE_RESP = 0xF
[10 empty] | Stream address | M2M2_ADDR_SENSOR_ADXL_STREAM = 0xC202
| Header | [8 empty] | Status | M2M2_APP_COMMON_STATUS_SUBSCRIBER_REMOVED = 0xA | Stream unsubscribed successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_SUBSCRIBER_COUNT_DECREMENT = 0xB | If stream subscribed more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 02 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0E; 2: 00; 3: 02; 4: C2
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 02 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0F; 2: 0A; 3: 02; 4: C2

## AdxlData
Sample ADXL data from CLI
RX::c2:02:c7:05:00:37:00:00:28:43:2f:1b:00:90:5e:db:01:f3:ff:23:00:33:01:82:02:f3:ff:23:00:33:01:82:02:f0:ff:26:00:34:01:81:02:e6:ff:26:00:38:01:82:02:f8:ff:1c:00:31:01
ADXL Data Stream | [8 empty] | source | M2M2_ADDR_SENSOR_ADXL_STREAM = 0xC202
[9 empty] | Destination | M2M2_ADDR_APP_CLI = 0xC705
[9 empty] | Command | M2M2_SENSOR_COMMON_CMD_STREAM_DATA = 0x28
Device to PC
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C2 | 02 | C7 | 05 | 00 | 37 | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13
Description | Command | Status | Sequence Number LSB | Sequence Number MSB | DATA Type
Data | 28 | 43 | 21 | 00 | 00
Payload
Byte Number | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
Description | Timestamp LSB | Timestamp | Timestamp | Timestamp MSB | 1st X Data LSB | 1st X Data MSB | 1st Y Data LSB | 1st Y Data MSB
Data | 00 | 82 | 1C | 03 | f3 | ff | 23 | 00
Payload
Byte Number | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29
Description | 1st Z Data LSB | 1st Z Data MSB | incTS LSB | incTS MSB | 2nd X Data LSB | 2nd X Data MSB | 2nd Y Data LSB | 2ndt Y Data MSB
Data | 33 | 01 | 82 | 02 | f3 | ff | 23 | 00
Payload
Byte Number | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37
Description | 2nd Z Data LSB | 2nd Z Data MSB | incTS LSB | incTS MSB | 3rd X Data LSB | 3rd X Data MSB | 3rd Y Data LSB | 3rd Y Data MSB
Data | 84 | 02 | CF | 03 | 00 | 00 | 00 | 00
Payload
Byte Number | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45
Description | 3rd Z Data LSB | 3rd Z Data MSB | incTS LSB | incTS MSB | 4th X Data LSB | 4th X Data MSB | 4th Y Data LSB | 4th Y Data MSB
Data | 84 | 02 | CF | 03 | 00 | 00 | 00 | 00
Payload
Byte Number | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53
Description | 4th Z Data LSB | 4th Z Data MSB | incTS LSB | incTS MSB | 5th X Data LSB | 5th X Data MSB | 5th Y Data LSB | 5th Y Data MSB
Data | 84 | 02 | CF | 03 | 00 | 00 | 00 | 00
Payload
Byte Number | 54 | 55
Description | 5th Z Data LSB | 5th Z Data MSB
Data | 84 | 02

## Reg_Rd_Wr
Register Read and Write Command
PC to Device
|  | Read ADPD register 0x08 | [7 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
Header | [9 empty] | Destination |  | M2M2_ADDR_SENSOR_ADPD = 0xC101
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | Request Command |  | M2M2_SENSOR_COMMON_CMD_READ_REG_16_REQ = 0x21 |  |  |  |  | Command to Perform a register operation on a device.
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB |  | Response Command |  | M2M2_SENSOR_COMMON_CMD_READ_REG_16_RESP = 0x22
Data | C7 | 05 | C1 | 01 | 00 | 0F | 00 | 00 |  | Status |  | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) |  |  |  |  | Register read/write operation performed successfully
Payload | [11 empty] | M2M2_APP_COMMON_STATUS_ERROR = 0x1 (or) |  |  |  |  | Error in Read/Write register opertaion
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 14
Description | Command | Status | Num_ops | Address LSB | Address MSB | Value LSB | Value MSB
Data | 21 | 00 | 01 | 08 | 00 | 00 | 00
Device to PC
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C1 | 01 | C7 | 05 | 00 | 0F | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 14
Description | Command | Status | Num_ops | Address LSB | Address MSB | Value LSB | Value MSB
Data | 22 | 00 | 01 | 08 | 00 | 08 | 16
PC to Device
Write ADPD register 0x34
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C7 | 05 | C1 | 01 | 00 | 0F | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 14
Description | Command | Status | Num_ops | Address LSB | Address MSB | Value LSB | Value MSB
Data | 23 | 00 | 01 | 34 | 00 | 00 | 03
Device to PC
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C1 | 01 | C7 | 05 | 00 | 0F | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 14
Description | Command | Status | Num_ops | Address LSB | Address MSB | Value LSB | Value MSB
Data | 24 | 00 | 01 | 34 | 00 | 00 | 03
PC to Device
|  | Multi Write ADPD register for 30 register addresses | [12 empty] | Note: To write 30 register add-value pairs, num_ops is made to be 30. With this total packet size would be (8 + 3 + 30*4) = 131 bytes. But you could choose num_ops, such that it would fit within the MTU of 244.
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C7 | 05 | C1 | 01 | 00 | 0F | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60 | 61 | 62 | 63 | 64 | 65 | 66 | 67 | 68 | 69 | 70 | 71 | 72 | 73 | 74 | 75 | 76 | 77 | 78 | 79 | 80 | 81 | 82 | 83 | 84 | 85 | 86 | 87 | 88 | 89 | 90 | 91 | 92 | 93 | 94 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 117 | 118 | 119 | 120 | 121 | 122 | 123 | 124 | 125 | 126 | 127 | 128 | 129 | 130 | 131
Description | Command | Status | Num_ops | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB
Data | 23 | 00 | 1E | 9 | 0 | 97 | 0 | 7 | 0 | ff | 8f | 0b | 0 | f6 | 2 | 0c | 0 | 12 | 0 | 0d | 0 | 10 | 27 | 0e | 0 | 0 | 0 | 0f | 0 | 6 | 0 | 14 | 0 | 0 | 80 | 20 | 0 | 22 | 22 | 21 | 0 | 0 | 0 | 22 | 0 | 3 | 0 | 23 | 0 | 2 | 3 | 24 | 0 | 1 | 0 | 31 | 1 | 0 | 0 | 51 | 1 | 0 | 0 | 71 | 1 | 0 | 0 | 91 | 1 | 0 | 0 | b1 | 1 | 0 | 0 | d1 | 1 | 0 | 0 | f1 | 1 | 0 | 0 | 11 | 2 | 0 | 0 | a0 | 1 | 0 | 40 | a1 | 1 | da | 41 | a2 | 1 | 5 | 0 | a3 | 1 | 2 | 50 | a4 | 1 | c0 | e3 | a5 | 1 | 5 | 0 | a6 | 1 | 0 | 0 | a7 | 1 | 40 | 1 | a8 | 1 | 0 | 0
PC to Device
Multi Write ADPD register for 30 register addresses RESP packet
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C7 | 05 | C1 | 01 | 00 | 0F | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60 | 61 | 62 | 63 | 64 | 65 | 66 | 67 | 68 | 69 | 70 | 71 | 72 | 73 | 74 | 75 | 76 | 77 | 78 | 79 | 80 | 81 | 82 | 83 | 84 | 85 | 86 | 87 | 88 | 89 | 90 | 91 | 92 | 93 | 94 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 117 | 118 | 119 | 120 | 121 | 122 | 123 | 124 | 125 | 126 | 127 | 128 | 129 | 130 | 131
Description | Command | Status | Num_ops | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB | Address LSB | Address MSB | Value LSB | Value MSB
Data | 24 | 00 | 1E | 9 | 0 | 97 | 0 | 7 | 0 | ff | 8f | 0b | 0 | f6 | 2 | 0c | 0 | 12 | 0 | 0d | 0 | 10 | 27 | 0e | 0 | 0 | 0 | 0f | 0 | 6 | 0 | 14 | 0 | 0 | 80 | 20 | 0 | 22 | 22 | 21 | 0 | 0 | 0 | 22 | 0 | 3 | 0 | 23 | 0 | 2 | 3 | 24 | 0 | 1 | 0 | 31 | 1 | 0 | 0 | 51 | 1 | 0 | 0 | 71 | 1 | 0 | 0 | 91 | 1 | 0 | 0 | b1 | 1 | 0 | 0 | d1 | 1 | 0 | 0 | f1 | 1 | 0 | 0 | 11 | 2 | 0 | 0 | a0 | 1 | 0 | 40 | a1 | 1 | da | 41 | a2 | 1 | 5 | 0 | a3 | 1 | 2 | 50 | a4 | 1 | c0 | e3 | a5 | 1 | 5 | 0 | a6 | 1 | 0 | 0 | a7 | 1 | 40 | 1 | a8 | 1 | 0 | 0
Read ADXL Register
PC to Device
|  | Read ADXL register 0x2D | [7 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
Header | [9 empty] | Destination |  | M2M2_ADDR_SENSOR_ADXL = 0xC102
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | Request Command |  | M2M2_SENSOR_COMMON_CMD_READ_REG_16_REQ = 0x21 |  |  |  |  | Command to Perform a register operation on a device.
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB |  | Response Command |  | M2M2_SENSOR_COMMON_CMD_READ_REG_16_RESP = 0x22
Data | C7 | 05 | C1 | 02 | 00 | 0F | 00 | 00 |  | Status |  | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) |  |  |  |  | Register read/write operation performed successfully
Payload | [11 empty] | M2M2_APP_COMMON_STATUS_ERROR = 0x1 (or) |  |  |  |  | Error in Read/Write register opertaion
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 14
Description | Command | Status | Num_ops | Address LSB | Address MSB | Value LSB | Value MSB
Data | 21 | 00 | 01 | 2d | 00 | 00 | 00
Device to PC
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C1 | 02 | C7 | 05 | 00 | 0F | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 14
Description | Command | Status | Num_ops | Address LSB | Address MSB | Value LSB | Value MSB
Data | 22 | 00 | 01 | 2d | 00 | 04 | 00
Write ADXL Register
PC to Device
|  | Write ADXL register 0x2C with value 0x82 | [7 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
Header | [9 empty] | Destination |  | M2M2_ADDR_SENSOR_ADXL = 0xC102
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | Request Command |  | M2M2_SENSOR_COMMON_CMD_WRITE_REG_16_REQ = 0x23 |  |  |  |  | Command to Perform a register operation on a device.
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB |  | Response Command |  | M2M2_SENSOR_COMMON_CMD_WRITE_REG_16_RESP = 0x24
Data | C7 | 05 | C1 | 02 | 00 | 0F | 00 | 00 |  | Status |  | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) |  |  |  |  | Register read/write operation performed successfully
Payload | [11 empty] | M2M2_APP_COMMON_STATUS_ERROR = 0x1 (or) |  |  |  |  | Error in Read/Write register opertaion
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 14
Description | Command | Status | Num_ops | Address LSB | Address MSB | Value LSB | Value MSB
Data | 23 | 00 | 01 | 2c | 00 | 82 | 00
Device to PC
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C1 | 02 | C7 | 05 | 00 | 0F | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 14
Description | Command | Status | Num_ops | Address LSB | Address MSB | Value LSB | Value MSB
Data | 24 | 00 | 01 | 2c | 00 | 82 | 00

## TemperatureStart
Temperature Start Commands
Start Temperature sensor | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_MED_TEMPERATURE = 0xC306
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_START_REQ = 0x4 | Command to start sensor
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_START_RESP = 0x5
[10 empty] | Status | M2M2_APP_COMMON_STATUS_STREAM_STARTED = 0x2 | Stream started successfully
| Header | [9 empty] | M2M2_APP_COMMON_STATUS_STREAM_IN_PROGRESS = 0x4 | Already Stream started
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_STREAM_NOT_STARTED = 0x7 | Error in starting stream
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 06 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 04; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 06 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 05; 2: 02; 3: 00; 4: 00
Subscribe Temperature Stream | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_MED_TEMPERATURE = 0xC306
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_SUBSCRIBE_REQ = 0xC | Command to subscribe the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_SUBSCRIBE_RESP = 0xD
[10 empty] | Stream address | M2M2_ADDR_MED_TEMPERATURE_STREAM = 0xC406
| Header | [8 empty] | Status | M2M2_APP_COMMON_STATUS_SUBSCRIBER_ADDED = 0x9 | Subscriber added successfully to the stream mentioned
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 06 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0C; 2: 00; 3: 06; 4: C4
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source LSB | Source MSB | Destination LSB | Destination MSB | Length LSB | Length MSB | Checksum LSB | Checksum MSB
| Data | C3 | 06 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0D; 2: 09; 3: 06; 4: C4

## TemperatureStop
Temperature Stop Commands
UnSubscribe Temperature Stream | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_MED_TEMPERATURE = 0xC306
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_UNSUBSCRIBE_REQ = 0xE | Command to unsubscribe the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_UNSUBSCRIBE_RESP = 0xF
[10 empty] | Stream address | M2M2_ADDR_MED_TEMPERATURE_STREAM = 0xC406
| Header | [8 empty] | Status | M2M2_APP_COMMON_STATUS_SUBSCRIBER_REMOVED = 0xA | Stream unsubscribed successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_SUBSCRIBER_COUNT_DECREMENT = 0xB | If stream subscribed more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 06 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0E; 2: 00; 3: 06; 4: C4
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 06 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0F; 2: 0A; 3: 06; 4: C4
Stop Temperature Sensor | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_MED_TEMPERATURE = 0xC306
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_STOP_REQ = 0x6 | Command to stop the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_STOP_RESP = 0x7
[10 empty] | Status | M2M2_APP_COMMON_STATUS_ERROR = 0x1
| Header | [9 empty] | M2M2_APP_COMMON_STATUS_STREAM_STOPPED = 0x3 | Stream stopped successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_STREAM_COUNT_DECREMENT =0x 6 | If stream started more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 06 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 06; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 06 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 07; 2: 03; 3: 00; 4: 00

## TemperatureData
Sample Temperature data from CLI
RX::c4:06:c7:05:00:14:00:00:28:43:09:04:da:e1:ff:01:60:81:70:81
Temperature Data Stream | [8 empty] | source | M2M2_ADDR_MED_TEMPERATURE_STREAM = 0xC406
[9 empty] | Destination | M2M2_ADDR_APP_CLI = 0xC705
[9 empty] | Command | M2M2_SENSOR_COMMON_CMD_STREAM_DATA = 0x28
Device to PC
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C4 | 06 | C7 | 05 | 00 | 14 | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16
Description | Command | Status | Sequence Number LSB | Sequence Number MSB | Timestamp LSB | Timestamp | Timestamp | Timestamp MSB
Data | 28 | 43 | 09 | 04 | DA | E1 | FF | 01
Payload
Byte Number | 17 | 18 | 19 | 20
Description | Temperature1 LSB | Temperature1 MSB | Temperature2 LSB | Temperature2 MSB
Data | 60 | 81 | 70 | 81
Temperature To Celsius Data Conversion
Example
Temperature1 LSB | Temperature1 MSB |  | Temperature1_d
60 | 81 |  | 33120
Temperature2 LSB | Temperature2 MSB |  | Temperature2_d
70 | 81 |  | 33136
|  |  | R1 = 39.0 R0 = 47.0 VDD = 2.5 fullScale = 65636 B = 4025 T0 = 25 + 273.16 vAdc = (VDD/fullScale)*Temperature1_d alpha = float(R1/R0) alpha = alpha*(vAdc/(VDD-vAdc)) Temperature1_celsius_value = B*T0/(B + math.log(alpha)*T0)-273.15 vAdc = (VDD/fullScale)*Temperature2_d alpha = float(R1/R0) alpha = alpha*(vAdc/(VDD-vAdc)) Temperature2_celsius_value = B*T0/(B + math.log(alpha)*T0)-273.15 | [5 empty] | vAdc = (2.5/65636)*33120 = 1.26150283381 alpha = float(39.0/47.0) = 0.829787234043 alpha = 0.829787234043 *(1.26150283381/(2.5- 1.26150283381)) = 0.845200922361 Temperature1_celsius_value = 4025*(25+273.16)/(4025 + math.log( 0.845200922361)*(25+273.16))-273.15 vAdc = (2.5/65636)*33136 = 1.26211225547 alpha = float(39.0/47.0) = 0.829787234043 alpha = 0.829787234043 *(1.26211225547/(2.5- 1.26211225547)) = 0.846025531915 Temperature2_celsius_value = 4025*(25+273.16)/(4025 + math.log( 0.846025531915)*(25+273.16))-273.15
Temperature1 in celsius = 28.7714398933
Temperature2 in celsius = 28.7493564452

## PPGMCStart
MC Stream Start Commands
Load Adpd configuration | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_SENSOR_ADPD = 0xC101
[13 empty] | Request Command |  | M2M2_SENSOR_ADPD_COMMAND_LOAD_CFG_REQ = 0x42 | Command to Load Adpd device configuration
| PC to Device | [11 empty] | Response Command |  | M2M2_SENSOR_ADPD_COMMAND_LOAD_CFG_RESP = 0x43
[13 empty] | Device ID |  | M2M2_SENSOR_ADPD_DEVICE_108 = 0x8
| Header | [11 empty] | Status |  | M2M2_APP_COMMON_STATUS_OK = 0 (or) | ADPD device loaded successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | [5 empty] | M2M2_APP_COMMON_STATUS_ERROR = 1 (or) | Error in Loading ADPD device
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 01 | 00 | 15 | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
| Description | Command | Status | Device ID | return data | return data | return data | return data | return data | return data | return data | return data | return data | return data
| Data | 42 | 00 | 08 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 01 | C7 | 05 | 00 | 15 | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
| Description | Command | Status | Device ID | return data | return data | return data | return data | return data | return data | return data | return data | return data | return data
| Data | 43 | 00 | 08 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00
Set PPG Library Configuration |  | 20: source | M2M2_ADDR_APP_CLI = 0xC705
|  | 20: Destination | M2M2_ADDR_MED_PPG = 0xC300
|  | 20: Request Command | M2M2_APP_COMMON_CMD_SET_LCFG_REQ = 0x14 | Command to Set the PPG LCFG configuration
| PC to Device | 20: Response Command | M2M2_APP_COMMON_CMD_SET_LCFG_RESP = 0x15
|  | 20: PPG LCFG ID | M2M2_SENSOR_PPG_LCFG_ID_ADPD108 = 0x6C
| Header | 20: Status | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) | PPG LCFG set successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | [5 empty] | M2M2_APP_COMMON_STATUS_ERROR = 1x0 (or) | Error in setting PPG LCFG
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 00 | 00 | 0B | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11
| Description | 1: Command; 2: Status; 3: PPG LCFG ID
| Data | 1: 14; 2: 00; 3: 6C
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 00 | C7 | 05 | 00 | 0A | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10
| Description | 1: Command; 2: Status
| Data | 1: 15; 2: 00
Clock Calibration | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_SENSOR_ADPD = 0xC101
[13 empty] | Request Command |  | M2M2_SENSOR_ADPD_COMMAND_CLOCK_CAL_REQ = 0x44 | Calibrate the 32M and 32K clock to reduce the deviation to a minimum
| PC to Device | [11 empty] | Response Command |  | M2M2_SENSOR_ADPD_COMMAND_CLOCK_CAL_RESP = 0x45
[13 empty] | Status |  | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) | Clock calibration done successfully
| Header | [13 empty] | M2M2_APP_COMMON_STATUS_ERROR = 0x1 (or) | Error in clock claibration
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 01 | 00 | 15 | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
| Description | Command | Status | Device ID | return data | return data | return data | return data | return data | return data | return data | return data | return data | return data
| Data | 44 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 01 | C7 | 05 | 00 | 15 | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
| Description | Command | Status | Device ID | return data | return data | return data | return data | return data | return data | return data | return data | return data | return data
| Data | 45 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00
Subscribe MC Stream |  | 20: source | M2M2_ADDR_APP_CLI = 0xC705
|  | 20: Destination | M2M2_ADDR_MED_PPG = 0xC300
|  | 20: Request Command | M2M2_APP_COMMON_CMD_STREAM_SUBSCRIBE_REQ = 0xC | Command to subscribe the stream
| PC to Device | 20: Response Command | M2M2_APP_COMMON_CMD_STREAM_SUBSCRIBE_RESP = 0xD
|  | 20: Stream address | M2M2_ADDR_SYS_PPGMC_STREAM = 0xC6D0
| Header | 20: Status | M2M2_APP_COMMON_STATUS_SUBSCRIBER_ADDED = 0x9 | Subscriber added successfully to the stream mentioned
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | c3 | 00 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0C; 2: 00; 3: D0; 4: C6
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 00 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0D; 2: 09; 3: D0; 4: C6
Start ADPD sensor | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_SENSOR_ADPD = 0xC101
[13 empty] | Request Command |  | M2M2_APP_COMMON_CMD_STREAM_START_REQ = 0x4 | Command to start sensor
| PC to Device | [11 empty] | Response Command |  | M2M2_APP_COMMON_CMD_STREAM_START_RESP = 0x5
[13 empty] | Status |  | M2M2_APP_COMMON_STATUS_STREAM_STARTED = 0x2 | Stream started successfully
| Header | [13 empty] | M2M2_APP_COMMON_STATUS_STREAM_IN_PROGRESS = 0x4 | Already Stream started
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | [5 empty] | M2M2_APP_COMMON_STATUS_STREAM_NOT_STARTED = 0x7 | Error in starting stream
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 01 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 04; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 01 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 05; 2: 02; 3: 00; 4: 00
Start ADXL sensor | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_SENSOR_ADXL = 0xC102
[13 empty] | Request Command |  | M2M2_APP_COMMON_CMD_STREAM_START_REQ = 0x4 | Command to start sensor
| PC to Device | [11 empty] | Response Command |  | M2M2_APP_COMMON_CMD_STREAM_START_RESP = 0x5
[13 empty] | Status |  | M2M2_APP_COMMON_STATUS_STREAM_STARTED = 0x2 | Stream started successfully
| Header | [13 empty] | M2M2_APP_COMMON_STATUS_STREAM_IN_PROGRESS = 0x4 | Already Stream started
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | [5 empty] | M2M2_APP_COMMON_STATUS_STREAM_NOT_STARTED = 0x7 | Error in starting stream
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 02 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 04; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source LSB | Source MSB | Destination LSB | Destination MSB | Length LSB | Length MSB | Checksum LSB | Checksum MSB
| Data | C1 | 02 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 05; 2: 02; 3: 00; 4: 00
Start PPG Application | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_MED_PPG = 0xC300
[13 empty] | Request Command |  | M2M2_APP_COMMON_CMD_STREAM_START_REQ = 0x4 | Command to start sensor
| PC to Device | [11 empty] | Response Command |  | M2M2_APP_COMMON_CMD_STREAM_START_RESP = 0x5
[13 empty] | Status |  | M2M2_APP_COMMON_STATUS_STREAM_STARTED = 0x2 | Stream started successfully
| Header | [13 empty] | M2M2_APP_COMMON_STATUS_STREAM_IN_PROGRESS = 0x4 | Already Stream started
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | [5 empty] | M2M2_APP_COMMON_STATUS_STREAM_NOT_STARTED = 0x7 | Error in starting stream
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 00 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 04; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 00 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 05; 2: 02; 3: 00; 4: 00

## PPGMCStop
MC Stream Stop Commands
Stop PPG Application | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_MED_PPG = 0xC300
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_STOP_REQ = 0x6 | Command to stop the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_STOP_RESP = 0x7
[10 empty] | Status | M2M2_APP_COMMON_STATUS_ERROR = 0x1
| Header | [9 empty] | M2M2_APP_COMMON_STATUS_STREAM_STOPPED = 0x3 | Stream stopped successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_STREAM_COUNT_DECREMENT =0x 6 | If stream started more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 00 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 06; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 00 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 07; 2: 03; 3: 00; 4: 00
UnSubscribe MC Stream | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_MED_PPG = 0xC300
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_UNSUBSCRIBE_REQ = 0xE | Command to unsubscribe the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_UNSUBSCRIBE_RESP = 0xF
[10 empty] | Stream address | M2M2_ADDR_SYS_PPGMC_STREAM = 0xC6D0
| Header | [8 empty] | Status | M2M2_APP_COMMON_STATUS_SUBSCRIBER_REMOVED = 0xA | Stream unsubscribed successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_SUBSCRIBER_COUNT_DECREMENT = 0xB | If stream subscribed more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 00 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0E; 2: 00; 3: D0; 4: C6
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 00 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0F; 2: 0A; 3: D0; 4: C6
Stop ADPD Sensor | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_SENSOR_ADPD = 0xC101
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_STOP_REQ = 0x6 | Command to stop the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_STOP_RESP = 0x7
[10 empty] | Status | M2M2_APP_COMMON_STATUS_ERROR = 0x1
| Header | [9 empty] | M2M2_APP_COMMON_STATUS_STREAM_STOPPED = 0x3 | Stream stopped successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_STREAM_COUNT_DECREMENT =0x 6 | If stream started more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 01 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 06; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source LSB | Source MSB | Destination LSB | Destination MSB | Length LSB | Length MSB | Checksum LSB | Checksum MSB
| Data | C1 | 01 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 07; 2: 03; 3: 00; 4: 00
Stop ADXL Sensor | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_SENSOR_ADXL = 0xC102
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_STOP_REQ = 0x6 | Command to stop the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_STOP_RESP = 0x7
[10 empty] | Status | M2M2_APP_COMMON_STATUS_ERROR = 0x1
| Header | [9 empty] | M2M2_APP_COMMON_STATUS_STREAM_STOPPED = 0x3 | Stream stopped successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_STREAM_COUNT_DECREMENT =0x 6 | If stream started more than one time
| Description | Source LSB | Source MSB | Destination LSB | Destination MSB | Length LSB | Length MSB | Checksum LSB | Checksum MSB
| Data | C7 | 05 | C1 | 02 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 06; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 02 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 07; 2: 03; 3: 00; 4: 00

## PPGMCData
Sample PPG-MC data from CLI
RX::c6:d0:c7:05:00:36:00:00:28:43:07:03:97:fe:55:06:00:49:ec:01:c5:46:00:00:71:02:00:cc:eb:01:92:3f:00:00:90:02:80:bf:eb:01:d4:38:00:00:71:02:80:1d:ed:01:10:14:00:00
MC Data Stream | [9 empty] | source |  | M2M2_ADDR_SYS_PPGMC_STREAM = 0xC6D0
[10 empty] | Destination |  | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Command |  | M2M2_SENSOR_COMMON_CMD_STREAM_DATA = 0x28
Device to PC
Header | [9 empty] | Note: | Sacled PPG signal and Motion compensated will be in 12.20 fixed point format and to get it in float format use below equation, ((float(x))/float(1048576)) where x - fixed ponit value
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C6 | D0 | C7 | 05 | 00 | 36 | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
Description | Command | Status | Sequence Number LSB | Sequence Number MSB | Timestamp LSB | Timestamp | Timestamp | Timestamp MSB | Scaled PPGSignal LSB | Scaled PPGSignal | Scaled PPGSignal | Scaled PPGSignal MSB
Data | 28 | 43 | 07 | 03 | 97 | FE | 55 | 06 | 00 | 49 | EC | 01
Payload
Byte Number | 21 | 22 | 23 | 24
Description | Motion Compensated LSB | Motion Compensated | Motion Compensated | Motion Compensated MSB
Data | C5 | 46 | 00 | 00
Payload
Byte Number | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34
Description | Timestamp LSB | Timestamp MSB | Scaled PPGSignal LSB | Scaled PPGSignal | Scaled PPGSignal | Scaled PPGSignal MSB | Motion Compensated LSB | Motion Compensated | Motion Compensated | Motion Compensated MSB
Data | 71 | 02 | 00 | CC | EB | 01 | 92 | 3F | 00 | 00
Payload
Byte Number | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44
Description | Timestamp LSB | Timestamp MSB | Scaled PPGSignal LSB | Scaled PPGSignal | Scaled PPGSignal | Scaled PPGSignal MSB | Motion Compensated LSB | Motion Compensated | Motion Compensated | Motion Compensated MSB
Data | 90 | 02 | 80 | BF | EB | 01 | D4 | 38 | 00 | 00
Payload
Byte Number | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54
Description | Timestamp LSB | Timestamp MSB | Scaled PPGSignal LSB | Scaled PPGSignal | Scaled PPGSignal | Scaled PPGSignal MSB | Motion Compensated LSB | Motion Compensated | Motion Compensated | Motion Compensated MSB
Data | 71 | 02 | 80 | 1D | ED | 01 | 10 | 14 | 00 | 00

## ECGStart
ECG Start Commands
Settting ECG Frequency
PC to Device
|  |  | Write ECG CFG 0x0 = 0x64 | [6 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
| Header | [8 empty] | Destination | M2M2_ADDR_MED_ECG = 0xC301
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | Request Command | M2M2_APP_COMMON_CMD_WRITE_LCFG_REQ = 0x18 |  | Command to Write the ECG LCFG configuration
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB | Response Command | M2M2_APP_COMMON_CMD_WRITE_LCFG_RESP = 0x19
| Data | C7 | 05 | C3 | 01 | 00 | 0E | 00 | 00 | Address, value | address = 0, value = 0x64
| Payload | [8 empty] | Status | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) |  | ECG LCFG written successfully
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 |  |  |  | M2M2_APP_COMMON_STATUS_ERROR = 1x0 (or) |  | Error in Writing ECG LCFG
| Description | Command | Status | Num_ops | Address | Value LSB | Value MSB
| Data | 18 | 00 | 01 | 00 | 64 | 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 01 | C7 | 05 | 00 | 0E | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14
| Description | Command | Status | Num_ops | Address | Value LSB | Value MSB
| Data | 19 | 00 | 01 | 00 | 64 | 00
Start ECG Sensor | [10 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[11 empty] | Destination | M2M2_ADDR_MED_ECG = 0xC301
[11 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_START_REQ = 0x4 |  | Command to start sensor
| PC to Device | [9 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_START_RESP = 0x5
[11 empty] | Status | M2M2_APP_COMMON_STATUS_STREAM_STARTED = 0x2 |  | Stream started successfully
| Header | [10 empty] | M2M2_APP_COMMON_STATUS_STREAM_IN_PROGRESS = 0x4 |  | Already Stream started
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  |  | M2M2_APP_COMMON_STATUS_STREAM_NOT_STARTED = 0x7 |  | Error in starting stream
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 01 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 04; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 01 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 05; 2: 02; 3: 00; 4: 00
Subscribe ECG Stream | [10 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[11 empty] | Destination | M2M2_ADDR_MED_ECG = 0xC301
[11 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_SUBSCRIBE_REQ = 0xC |  | Command to subscribe the stream
| PC to Device | [9 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_SUBSCRIBE_RESP = 0xD
[11 empty] | Stream address | M2M2_ADDR_MED_ECG_STREAM = 0xC401
| Header | [9 empty] | Status | M2M2_APP_COMMON_STATUS_SUBSCRIBER_ADDED = 0x9 |  | Subscriber added successfully to the stream mentioned
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 01 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0C; 2: 00; 3: 01; 4: C4
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 01 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0D; 2: 09; 3: 01; 4: C4

## ECGStop
ECG Stop Commands
UnSubscribe ECG Stream | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_MED_ECG = 0xC301
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_UNSUBSCRIBE_REQ = 0xE | Command to unsubscribe the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_UNSUBSCRIBE_RESP = 0xF
[10 empty] | Stream address | M2M2_ADDR_MED_ECG_STREAM = 0xC401
| Header | [8 empty] | Status | M2M2_APP_COMMON_STATUS_SUBSCRIBER_REMOVED = 0xA | Stream unsubscribed successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_SUBSCRIBER_COUNT_DECREMENT = 0xB | If stream subscribed more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 01 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0E; 2: 00; 3: 01; 4: C4
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 01 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0F; 2: 0A; 3: 01; 4: C4
Stop ECG Sensor | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_MED_ECG = 0xC301
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_STOP_REQ = 0x6 | Command to stop the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_STOP_RESP = 0x7
[10 empty] | Status | M2M2_APP_COMMON_STATUS_ERROR = 0x1
| Header | [9 empty] | M2M2_APP_COMMON_STATUS_STREAM_STOPPED = 0x3 | Stream stopped successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_STREAM_COUNT_DECREMENT =0x 6 | If stream started more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 01 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 06; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 01 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 07; 2: 03; 3: 00; 4: 00

## ECGData
Sample ECG data from CLI
c4:01:c7:05:00:3d:00:00:28:43:07:00:01:f5:5f:1b:19:01:77:fc:ff:39:01:00:00:38:01:fc:ff:39:01:00:00:38:01:fb:ff:39:01:00:00:38:01:60:e7:40:01:00:00:41:01:47:d3:40:01:cb:00:40:01:92:a6
ECG Data Stream | [8 empty] | source | M2M2_ADDR_MED_ECG_STREAM = 0xC401
[9 empty] | Destination | M2M2_ADDR_APP_CLI = 0xC705
[9 empty] | Command | M2M2_SENSOR_COMMON_CMD_STREAM_DATA = 0x28
Device to PC
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C4 | 01 | C7 | 05 | 00 | 3D | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13
Description | Command | Status | Sequence Number LSB | Sequence Number MSB | DATA Type
Data | 28 | 43 | 07 | 00 | 01
Payload
Byte Number | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
Description | Timestamp LSB | Timestamp | Timestamp | Timestamp MSB | ECG info (Eletrode touch status) | HR | First ECG data LSB | First ECG Data MSB
Data | F5 | 5F | 1B | 19 | 01 | 77 | FC | FF
Payload
Byte Number | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29
Description | Timestamp LSB | Timestamp MSB | ECG DATA LSB | ECG DATA MSB | Timestamp LSB | Timestamp MSB | ECG DATA LSB | ECG DATA MSB
Data | 39 | 01 | 00 | 00 | 38 | 01 | FC | FF
Payload
Byte Number | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37
Description | Timestamp LSB | Timestamp MSB | ECG DATA LSB | ECG DATA MSB | Timestamp LSB | Timestamp MSB | ECG DATA LSB | ECG DATA MSB
Data | 39 | 01 | 00 | 00 | 38 | 01 | FB | FF
Payload
Byte Number | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45
Description | Timestamp LSB | Timestamp MSB | ECG DATA LSB | ECG DATA MSB | Timestamp LSB | Timestamp MSB | ECG DATA LSB | ECG DATA MSB
Data | 39 | 01 | 00 | 00 | 38 | 01 | 60 | E7
Payload
Byte Number | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53
Description | Timestamp LSB | Timestamp MSB | ECG DATA LSB | ECG DATA MSB | Timestamp LSB | Timestamp MSB | ECG DATA LSB | ECG DATA MSB
Data | 40 | 01 | 00 | 00 | 41 | 01 | 47 | D3
Payload
Byte Number | 54 | 55 | 56 | 57 | 58 | 59 | 60 | 61
Description | Timestamp LSB | Timestamp MSB | ECG DATA LSB | ECG DATA MSB | Timestamp LSB | Timestamp MSB | ECG DATA LSB | ECG DATA MSB
Data | 40 | 01 | CB | 00 | 40 | 01 | 92 | A6

## PedometerStart
Pedometer Start Commands
Start ADXL sensor | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_SENSOR_ADXL = 0xC102
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_START_REQ = 0x4 | Command to start sensor
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_START_RESP = 0x5
[10 empty] | Status | M2M2_APP_COMMON_STATUS_STREAM_STARTED = 0x2 | Stream started successfully
| Header | [9 empty] | M2M2_APP_COMMON_STATUS_STREAM_IN_PROGRESS = 0x4 | Already Stream started
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_STREAM_NOT_STARTED = 0x7 | Error in starting stream
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 02 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 04; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 02 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 05; 2: 02; 3: 00; 4: 00
Start Pedometer | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_MED_PED = 0xC304
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_START_REQ = 0x4 | Command to start sensor
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_START_RESP = 0x5
[10 empty] | Status | M2M2_APP_COMMON_STATUS_STREAM_STARTED = 0x2 | Stream started successfully
| Header | [9 empty] | M2M2_APP_COMMON_STATUS_STREAM_IN_PROGRESS = 0x4 | Already Stream started
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_STREAM_NOT_STARTED = 0x7 | Error in starting stream
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 04 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 04; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source LSB | Source MSB | Destination LSB | Destination MSB | Length LSB | Length MSB | Checksum LSB | Checksum MSB
| Data | C3 | 04 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 05; 2: 02; 3: 00; 4: 00
Subscribe Pedometer Stream | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_MED_PED = 0xC304
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_SUBSCRIBE_REQ = 0xC | Command to subscribe the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_SUBSCRIBE_RESP = 0xD
[10 empty] | Stream address | M2M2_ADDR_MED_PED_STREAM = 0xC404
| Header | [8 empty] | Status | M2M2_APP_COMMON_STATUS_SUBSCRIBER_ADDED = 0x9 | Subscriber added successfully to the stream mentioned
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 04 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0C; 2: 00; 3: 04; 4: C4
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 04 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0D; 2: 09; 3: 04; 4: C4

## PedometerStop
Pedometer Stop Commands
Stop Pedometer | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_MED_PED = 0xC304
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_STOP_REQ = 0x6 | Command to stop the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_STOP_RESP = 0x7
[10 empty] | Status | M2M2_APP_COMMON_STATUS_ERROR = 0x1
| Header | [9 empty] | M2M2_APP_COMMON_STATUS_STREAM_STOPPED = 0x3 | Stream stopped successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_STREAM_COUNT_DECREMENT =0x 6 | If stream started more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 04 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 06; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 04 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 07; 2: 03; 3: 00; 4: 00
UnSubscribe Pedometer Stream | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_MED_PED = 0xC304
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_UNSUBSCRIBE_REQ = 0xE | Command to unsubscribe the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_UNSUBSCRIBE_RESP = 0xF
[10 empty] | Stream address | M2M2_ADDR_MED_PED_STREAM = 0xC404
| Header | [8 empty] | Status | M2M2_APP_COMMON_STATUS_SUBSCRIBER_REMOVED = 0xA | Stream unsubscribed successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_SUBSCRIBER_COUNT_DECREMENT = 0xB | If stream subscribed more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 04 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0E; 2: 00; 3: 04; 4: C4
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 04 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0F; 2: 0A; 3: 04; 4: C4
Stop ADXL Sensor | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_SENSOR_ADXL = 0xC102
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_STOP_REQ = 0x6 | Command to stop the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_STOP_RESP = 0x7
[10 empty] | Status | M2M2_APP_COMMON_STATUS_ERROR = 0x1
| Header | [9 empty] | M2M2_APP_COMMON_STATUS_STREAM_STOPPED = 0x3 | Stream stopped successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_STREAM_COUNT_DECREMENT =0x 6 | If stream started more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 02 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 06; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 02 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 07; 2: 03; 3: 00; 4: 00

## PedometerData
Sample PEDOMETER data from CLI
c4:04:c7:05:00:17:00:00:28:43:63:00:00:00:00:00:00:00:95:23:f8:19:00
PEDOMETER Data Stream | [9 empty] | source | M2M2_ADDR_MED_PED_STREAM = 0xC404
[10 empty] | Destination | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Command | M2M2_SENSOR_COMMON_CMD_STREAM_DATA = 0x28
Device to PC
Header
Byte Number | 1 | 2 | 3 | 4 |  | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C4 | 04 | C7 | 05 | 00 | 17 | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18
Description | Command | Status | Sequence Number LSB | Sequence Number MSB | Steps LSB | Steps | Steps | Steps MSB | Algo Status LSB | Algo Status MSB
Data | 28 | 43 | 63 | 00 | 00 | 00 | 00 | 00 | 00 | 00
Payload
Byte Number | 19 | 20 | 21 | 22 | 23
Description | TimeStamp LSB | TimeStamp | TimeStamp | TimeStamp MSB | Reserved
Data | 95 | 23 | F8 | 19 | 00

## FileSystem_Commands
File System Commands
File system format command | [12 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination | M2M2_ADDR_SYS_FS= 0xC501
[13 empty] | Request Command | M2M2_FILE_SYS_CMD_FORMAT_REQ = 0x46 | Command to format the file system
| PC to Device | [11 empty] | Response Command | M2M2_FILE_SYS_CMD_FORMAT_RESP = 0x47
[13 empty] | Status | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) | File system formatted successfully performed successfully
| Header | [12 empty] | M2M2_FILE_SYS_ERR_FORMAT = 0x47 | Error in formatting the file system
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C5 | 01 | 00 | 0A | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10
| Description | 1: Command; 2: Status
| Data | 1: 46; 2: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | 01 | C5 | 05 | C7 | 0A | 00 | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10
| Description | 1: Command; 2: Status
| Data | 1: 47; 2: 00
File open Command | [12 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination | M2M2_ADDR_SYS_FS= 0xC501
[13 empty] | Request Command | M2M2_FILE_SYS_CMD_START_LOGGING_REQ = 0x76 | Command to open new file
| PC to Device | [11 empty] | Response Command | M2M2_FILE_SYS_CMD_START_LOGGING_RESP = 0x77
[13 empty] | Status | M2M2_FILE_SYS_STATUS_OK = 0x41 (or) | File created successfully performed successfully
| Header | [12 empty] | M2M2_FILE_SYS_STATUS_LOGGING_IN_PROGRESS = 0x4E | Already File in open state, Close the existing file to create now one
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  |  |  |  | M2M2_FILE_SYS_STATUS_ERROR = 0x42 | Error in Creating file
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB |  |  |  |  | M2M2_FILE_SYS_ERR_MAX_FILE_COUNT = 0x4A | Maximum file count reached(64), Format the file system to create new file.
| Data | C7 | 05 | C5 | 01 | 00 | 0C | 00 | 00 |  |  |  |  | M2M2_FILE_SYS_ERR_INVALID = 0x45 | Invalid file directory
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 76; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 01 | C7 | 05 | 00 | 0A | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10
| Description | 1: Command; 2: Status
| Data | 1: 77; 2: 41
File Close Command | [12 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination | M2M2_ADDR_SYS_FS= 0xC501
[13 empty] | Request Command | M2M2_FILE_SYS_CMD_STOP_LOGGING_REQ = 0x78 | Command to close the file
| PC to Device | [11 empty] | Response Command | M2M2_FILE_SYS_CMD_STOP_LOGGING_RESP = 0x79
[13 empty] | Status | M2M2_FILE_SYS_STATUS_LOGGING_STOPPED = 0x4D (or) | File closed successfully performed successfully
| Header | [12 empty] | M2M2_FILE_SYS_STATUS_ERROR = 0x42 (or) | Error in Closing the file
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  |  |  |  | M2M2_FILE_SYS_STATUS_LOGGING_NOT_STARTED = 0x50 | Logging not yet started
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C5 | 01 | 00 | 0B | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11
| Description | 1: Command; 2: Status; 3: stop type
| Data | 1: 78; 2: 00; 3: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 01 | C7 | 05 | 00 | 0B | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11
| Description | 1: Command; 2: Status; 3: stop type
| Data | 1: 79; 2: 41; 3: 00
File Staus command
[13 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination | M2M2_ADDR_SYS_FS= 0xC501
| PC to Device | [11 empty] | Request Command | M2M2_FILE_SYS_CMD_GET_FS_STATUS_REQ = 0x50 | Command to close the file
[13 empty] | Response Command | M2M2_FILE_SYS_CMD_GET_FS_STATUS_RESP = 0x51
| Header | [11 empty] | Status | M2M2_APP_COMMON_STATUS_STREAM_IN_PROGRESS = 0x4 | File is in open state performed successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  |  |  |  | M2M2_APP_COMMON_STATUS_STREAM_NOT_STARTED = 0x7 | No file has been opened
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C5 | 01 | 00 | 0A | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10
| Description | 1: Command; 2: Status
| Data | 1: 50; 2: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 01 | C7 | 05 | 00 | 0A | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10
| Description | 1: Command; 2: Status
| Data | 1: 51; 2: 07
File System volume info command
[13 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination | M2M2_ADDR_SYS_FS= 0xC501
| PC to Device | [11 empty] | Request Command | M2M2_FILE_SYS_CMD_VOL_INFO_REQ = 0x4E | Command to get the volume info of the file system
[13 empty] | Response Command | M2M2_FILE_SYS_CMD_VOL_INFO_RESP = 0x4F
| Header | [11 empty] | Status | M2M2_FILE_SYS_STATUS_OK = 0x41 | Volume info obtained successfully performed successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  |  |  |  | M2M2_FILE_SYS_STATUS_ERROR = 0x42 | Error in getting volume info
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C5 | 01 | 00 | 0A | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10
| Description | 1: Command; 2: Status
| Data | 1: 4E; 2: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 01 | C7 | 05 | 00 | 14 | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
| Description | Command | Status | Total memory in bytes LSB | Total memory in bytes | Total memory in bytes | Total memory in bytes MSB | Used Memory in bytes LSB | Used Memory in bytes | Used Memory in bytes | Used Memory in bytes MSB | Free Memory in percentage LSB | Free Memory in percentage MSB
| Data | 4F | 41 | 00 | 00 | 00 | 20 | 00 | 00 | 0A | 00 | 63 | 00
FS Stream Subscribe |  | 20: source | M2M2_ADDR_APP_CLI = 0xC705
|  | 20: Destination | M2M2_ADDR_SYS_FS= 0xC501
|  | 20: Request Command | M2M2_FILE_SYS_CMD_LOG_STREAM_REQ = 0x72 | Command to subscribe the stream
| PC to Device | 20: Response Command | M2M2_FILE_SYS_CMD_LOG_STREAM_RESP = 0x73
|  | 20: Stream address | M2M2_ADDR_MED_ECG_STREAM = 0xC401
| Header | 20: Status | M2M2_APP_COMMON_STATUS_SUBSCRIBER_ADDED = 0x9 | Subscriber added successfully to the stream mentioned
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C5 | 01 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 72; 2: 00; 3: 01; 4: C4
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 01 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 73; 2: 09; 3: 01; 4: C4
FS Stream UnSubscribe | [12 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination | M2M2_ADDR_SYS_FS= 0xC501
[13 empty] | Request Command | M2M2_FILE_SYS_CMD_STOP_STREAM_REQ = 0x74 | Command to unsubscribe the stream
| PC to Device | [11 empty] | Response Command | M2M2_FILE_SYS_CMD_STOP_STREAM_REQ = 0x75
[13 empty] | Stream address | M2M2_ADDR_MED_ECG_STREAM = 0xC401
| Header | [11 empty] | Status | M2M2_APP_COMMON_STATUS_SUBSCRIBER_REMOVED = 0xA | Stream unsubscribed successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  |  |  |  | M2M2_APP_COMMON_STATUS_SUBSCRIBER_COUNT_DECREMENT = 0xB | If stream subscribed more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C5 | 01 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 74; 2: 00; 3: 01; 4: C4
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 01 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 75; 2: 0A; 3: 01; 4: C4
FS Stream sub/unsub Status | [12 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination | M2M2_ADDR_SYS_FS= 0xC501
[13 empty] | Request Command | M2M2_FILE_SYS_CMD_GET_FS_STREAM_SUB_STATUS_REQ= 0x52 | Command to get the state of stream subscribed to FS
| PC to Device | [11 empty] | Response Command | M2M2_FILE_SYS_CMD_GET_FS_STREAM_SUB_STATUS_RESP = 0x53
[13 empty] | Stream address | M2M2_ADDR_MED_ECG_STREAM = 0xC401
| Header | [11 empty] | Status | M2M2_FILE_SYS_STATUS_OK = 0x41 | Subscription state obtained successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  |  |  |  | M2M2_FILE_SYS_STATUS_ERROR = 0x42 | Error in getting Subscription state
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB |  |  |  | Subscribed state | M2M2_FILE_SYS_UNSUBSCRIBED = 0x0 M2M2_FILE_SYS_SUBSCRIBED = 0x1 | Subscribe State of stream
| Data | C7 | 05 | C5 | 01 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 52; 2: 00; 3: 01; 4: C4
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 01 | C7 | 05 | 00 | 0D | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13
| Description | Command | Status | Stream Address LSB | Stream Address MSB | Subscribe State
| Data | 53 | 41 | 01 | C4 | 01

## PPG_logging_start
PPG Logging Start Commands
FS Subscribe ADPD Stream | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_SYS_FS= 0xC501
[13 empty] | Request Command |  | M2M2_FILE_SYS_CMD_LOG_STREAM_REQ = 0x72 | Command to subscribe the stream
| PC to Device | [11 empty] | Response Command |  | M2M2_FILE_SYS_CMD_LOG_STREAM_RESP = 0x73
[13 empty] | Stream address |  | M2M2_ADDR_MED_PPG_STREAM = 0xC400
| Header | [11 empty] | Status |  | M2M2_APP_COMMON_STATUS_SUBSCRIBER_ADDED = 0x9 | Subscriber added successfully to the stream mentioned
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C5 | 01 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 72; 2: 00; 3: 00; 4: C4
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 01 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 73; 2: 09; 3: 00; 4: C4
Load Adpd configuration | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_SENSOR_ADPD4000 = 0xC110
[13 empty] | Request Command |  | M2M2_SENSOR_ADPD_COMMAND_LOAD_CFG_REQ = 0x42 | Command to Load Adpd device configuration
| PC to Device | [11 empty] | Response Command |  | M2M2_SENSOR_ADPD_COMMAND_LOAD_CFG_RESP = 0x43
[13 empty] | Device ID |  | M2M2_SENSOR_ADPD4000_DEVICE_4000_G = 0x28
| Header | [11 empty] | Status |  | M2M2_APP_COMMON_STATUS_OK = 0 (or) | ADPD device loaded successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | [5 empty] | M2M2_APP_COMMON_STATUS_ERROR = 1 (or) | Error in Loading ADPD device
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 0 | 15 | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
| Description | Command | Status | Device ID | return data | return data | return data | return data | return data | return data | return data | return data | return data | return data
| Data | 42 | 00 | 28 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 10 | C7 | 05 | 00 | 15 | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
| Description | Command | Status | Device ID | return data | return data | return data | return data | return data | return data | return data | return data | return data | return data
| Data | 43 | 00 | 28 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00
Clock Calibration |  | 20: source | M2M2_ADDR_APP_CLI = 0xC705
|  | 20: Destination | M2M2_ADDR_SENSOR_ADPD4000 = 0xC110
|  | 20: Request Command | M2M2_SENSOR_ADPD_COMMAND_CLOCK_CAL_REQ = 0x44 | Calibrate the 32M and 32K clock to reduce the deviation to a minimum
| PC to Device | 20: Response Command | M2M2_SENSOR_ADPD_COMMAND_CLOCK_CAL_RESP = 0x45
|  | 20: Status | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) | Clock calibration done successfully
| Header | [13 empty] | M2M2_APP_COMMON_STATUS_ERROR = 0x1 (or) | Error in clock claibration
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  |  |  | Device ID |  | 6-> in DVT1 Watches; 2 -> DVT2 Watches
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 00 | 08 | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11
| Description | 1: Command; 2: Status; 3: Device ID
| Data | 1: 44; 2: 00; 3: 06
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 10 | C7 | 05 | 00 | 08 | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11
| Description | 1: Command; 2: Status; 3: Device ID
| Data | 1: 45; 2: 00; 3: 06
Set PPG Library Configuration | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_MED_PPG = 0xC300
[13 empty] | Request Command |  | M2M2_APP_COMMON_CMD_SET_LCFG_REQ = 0x14 | Command to Set the PPG LCFG configuration
| PC to Device | [11 empty] | Response Command |  | M2M2_APP_COMMON_CMD_SET_LCFG_RESP = 0x15
[13 empty] | PPG LCFG ID |  | M2M2_SENSOR_PPG_LCFG_ID_ADPD4000 = 0x28
| Header | [11 empty] | Status |  | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) | PPG LCFG set successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | [5 empty] | M2M2_APP_COMMON_STATUS_ERROR = 1x0 (or) | Error in setting PPG LCFG
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 00 | 00 | 08 | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11
| Description | 1: Command; 2: Status; 3: PPG LCFG ID
| Data | 1: 14; 2: 00; 3: 28
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 00 | C7 | 05 | 00 | 0A | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10
| Description | 1: Command; 2: Status
| Data | 1: 15; 2: 00
Start PPG Application | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_MED_PPG = 0xC300
[13 empty] | Request Command |  | M2M2_APP_COMMON_CMD_STREAM_START_REQ = 0x4 | Command to start sensor
| PC to Device | [11 empty] | Response Command |  | M2M2_APP_COMMON_CMD_STREAM_START_RESP = 0x5
[13 empty] | Status |  | M2M2_APP_COMMON_STATUS_STREAM_STARTED = 0x2 | Stream started successfully
| Header | [13 empty] | M2M2_APP_COMMON_STATUS_STREAM_IN_PROGRESS = 0x4 | Already Stream started
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | [5 empty] | M2M2_APP_COMMON_STATUS_STREAM_NOT_STARTED = 0x7 | Error in starting stream
| Description | Source LSB | Source MSB | Destination LSB | Destination MSB | Length LSB | Length MSB | Checksum LSB | Checksum MSB
| Data | C7 | 05 | C3 | 00 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 04; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source LSB | Source MSB | Destination LSB | Destination MSB | Length LSB | Length MSB | Checksum LSB | Checksum MSB
| Data | C3 | 00 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 05; 2: 02; 3: 00; 4: 00
File open Command | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_SYS_FS= 0xC501
[13 empty] | Request Command |  | M2M2_FILE_SYS_CMD_START_LOGGING_REQ = 0x76 | Command to open new file
| PC to Device | [11 empty] | Response Command |  | M2M2_FILE_SYS_CMD_START_LOGGING_RESP = 0x77
[13 empty] | Status |  | M2M2_FILE_SYS_STATUS_OK = 0x41 (or) | File created successfully performed successfully
| Header | [13 empty] | M2M2_FILE_SYS_STATUS_LOGGING_IN_PROGRESS = 0x4E | Already File in open state, Close the existing file to create now one
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | [5 empty] | M2M2_FILE_SYS_STATUS_ERROR = 0x42 | Error in Creating file
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB | [5 empty] | M2M2_FILE_SYS_ERR_MAX_FILE_COUNT = 0x4A | Maximum file count reached(64), Format the file system to create new file.
| Data | C7 | 05 | C5 | 01 | 00 | 0C | 00 | 00 | [5 empty] | M2M2_FILE_SYS_ERR_INVALID = 0x45 | Invalid file directory
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 76; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 01 | C7 | 05 | 00 | 0A | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10
| Description | 1: Command; 2: Status
| Data | 1: 77; 2: 41

## Ppg_logging_stop
PPG Logging Stop Commands
Stop PPG Application | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_MED_PPG = 0xC300
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_STOP_REQ = 0x6 | Command to stop the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_STOP_RESP = 0x7
[10 empty] | Status | M2M2_APP_COMMON_STATUS_ERROR = 0x1
| Header | [9 empty] | M2M2_APP_COMMON_STATUS_STREAM_STOPPED = 0x3 | Stream stopped successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_STREAM_COUNT_DECREMENT =0x 6 | If stream started more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 00 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 06; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 00 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 07; 2: 03; 3: 00; 4: 00
UnSubscribe FS PPG Stream | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_SYS_FS= 0xC501
[10 empty] | Request Command | M2M2_FILE_SYS_CMD_STOP_STREAM_REQ = 0x74 | Command to unsubscribe the stream
| PC to Device | [8 empty] | Response Command | M2M2_FILE_SYS_CMD_STOP_STREAM_REQ = 0x75
[10 empty] | Stream address | M2M2_ADDR_MED_PPG_STREAM = 0xC400
| Header | [8 empty] | Status | M2M2_APP_COMMON_STATUS_SUBSCRIBER_REMOVED = 0xA | Stream unsubscribed successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_SUBSCRIBER_COUNT_DECREMENT = 0xB | If stream subscribed more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C5 | 01 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 74; 2: 00; 3: 00; 4: C4
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 01 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 75; 2: 0A; 3: 00; 4: C4
File Close Command | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_SYS_FS= 0xC501
[10 empty] | Request Command | M2M2_FILE_SYS_CMD_STOP_LOGGING_REQ = 0x78 | Command to close the file
| PC to Device | [8 empty] | Response Command | M2M2_FILE_SYS_CMD_STOP_LOGGING_RESP = 0x79
[10 empty] | Status | M2M2_FILE_SYS_STATUS_LOGGING_STOPPED = 0x4D (or) | File closed successfully performed successfully
| Header | [9 empty] | M2M2_FILE_SYS_STATUS_ERROR = 0x42 (or) | Error in Closing the file
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_FILE_SYS_STATUS_LOGGING_NOT_STARTED = 0x50 | Logging not yet started
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C5 | 01 | 00 | 0B | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11
| Description | 1: Command; 2: Status; 3: stop type
| Data | 1: 78; 2: 00; 3: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 01 | C7 | 05 | 00 | 0B | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11
| Description | 1: Command; 2: Status; 3: stop type
| Data | 1: 79; 2: 41; 3: 00

## ECG_logging_start
ECG Start Commands
FS Subscribe ECG Stream | [10 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[11 empty] | Destination | M2M2_ADDR_SYS_FS= 0xC501
[11 empty] | Request Command | M2M2_FILE_SYS_CMD_LOG_STREAM_REQ = 0x72 |  | Command to subscribe the stream
| PC to Device | [9 empty] | Response Command | M2M2_FILE_SYS_CMD_LOG_STREAM_RESP = 0x73
[11 empty] | Stream address | M2M2_ADDR_MED_ECG_STREAM = 0xC401
| Header | [9 empty] | Status | M2M2_APP_COMMON_STATUS_SUBSCRIBER_ADDED = 0x9 |  | Subscriber added successfully to the stream mentioned
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C5 | 01 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 72; 2: 00; 3: 01; 4: C4
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 01 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 73; 2: 09; 3: 01; 4: C4
Settting ECG Frequency
PC to Device
|  | 2: Write ECG CFG 0x0 = 0x64 | source | M2M2_ADDR_APP_CLI = 0xC705
| Header | [8 empty] | Destination | M2M2_ADDR_MED_ECG = 0xC301
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | Request Command | M2M2_APP_COMMON_CMD_WRITE_LCFG_REQ = 0x18 |  | Command to Write the ECG LCFG configuration
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB | Response Command | M2M2_APP_COMMON_CMD_WRITE_LCFG_RESP = 0x19
| Data | C7 | 05 | C3 | 1 | 00 | 0E | 00 | 00 | Address, value | address = 0, value = 0x64
| Payload | [8 empty] | Status | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) |  | ECG LCFG written successfully
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 |  |  |  | M2M2_APP_COMMON_STATUS_ERROR = 1x0 (or) |  | Error in Writing ECG LCFG
| Description | Command | Status | Num_ops | Address | Value LSB | Value MSB
| Data | 18 | 00 | 01 | 00 | 64 | 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 1 | C7 | 05 | 00 | 0E | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14
| Description | Command | Status | Num_ops | Address | Value LSB | Value MSB
| Data | 19 | 00 | 01 | 00 | 64 | 00
Start ECG Sensor | [10 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[11 empty] | Destination | M2M2_ADDR_MED_ECG = 0xC301
[11 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_START_REQ = 0x4 |  | Command to start sensor
| PC to Device | [9 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_START_RESP = 0x5
[11 empty] | Status | M2M2_APP_COMMON_STATUS_STREAM_STARTED = 0x2 |  | Stream started successfully
| Header | [10 empty] | M2M2_APP_COMMON_STATUS_STREAM_IN_PROGRESS = 0x4 |  | Already Stream started
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  |  | M2M2_APP_COMMON_STATUS_STREAM_NOT_STARTED = 0x7 |  | Error in starting stream
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 01 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 04; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 01 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 05; 2: 02; 3: 00; 4: 00
File open Command | [10 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[11 empty] | Destination | M2M2_ADDR_SYS_FS= 0xC501
[11 empty] | Request Command | M2M2_FILE_SYS_CMD_START_LOGGING_REQ = 0x76 |  | Command to open new file
| PC to Device | [9 empty] | Response Command | M2M2_FILE_SYS_CMD_START_LOGGING_RESP = 0x77
[11 empty] | Status | M2M2_FILE_SYS_STATUS_OK = 0x41 (or) |  | File created successfully performed successfully
| Header | [10 empty] | M2M2_FILE_SYS_STATUS_LOGGING_IN_PROGRESS = 0x4E |  | Already File in open state, Close the existing file to create now one
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  |  | M2M2_FILE_SYS_STATUS_ERROR = 0x42 |  | Error in Creating file
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB |  |  | M2M2_FILE_SYS_ERR_MAX_FILE_COUNT = 0x4A |  | Maximum file count reached(64), Format the file system to create new file.
| Data | C7 | 05 | C5 | 01 | 00 | 0C | 00 | 00 |  |  | M2M2_FILE_SYS_ERR_INVALID = 0x45 |  | Invalid file directory
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 76; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 01 | C7 | 05 | 00 | 0A | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10
| Description | 1: Command; 2: Status
| Data | 1: 77; 2: 41

## ECG_Logging_Stop
ECG Stop Commands
Stop ECG Sensor | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_MED_ECG = 0xC301
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_STOP_REQ = 0x6 | Command to stop the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_STOP_RESP = 0x7
[10 empty] | Status | M2M2_APP_COMMON_STATUS_ERROR = 0x1
| Header | [9 empty] | M2M2_APP_COMMON_STATUS_STREAM_STOPPED = 0x3 | Stream stopped successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_STREAM_COUNT_DECREMENT =0x 6 | If stream started more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C3 | 01 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 06; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C3 | 01 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 07; 2: 03; 3: 00; 4: 00
UnSubscribe FS ECG Stream | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_SYS_FS= 0xC501
[10 empty] | Request Command | M2M2_FILE_SYS_CMD_STOP_STREAM_REQ = 0x74 | Command to unsubscribe the stream
| PC to Device | [8 empty] | Response Command | M2M2_FILE_SYS_CMD_STOP_STREAM_REQ = 0x75
[10 empty] | Stream address | M2M2_ADDR_MED_ECG_STREAM = 0xC401
| Header | [8 empty] | Status | M2M2_APP_COMMON_STATUS_SUBSCRIBER_REMOVED = 0xA | Stream unsubscribed successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_SUBSCRIBER_COUNT_DECREMENT = 0xB | If stream subscribed more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C5 | 01 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 74; 2: 00; 3: 01; 4: C4
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 01 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 75; 2: 0A; 3: 01; 4: C4
File Close Command | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_SYS_FS= 0xC501
[10 empty] | Request Command | M2M2_FILE_SYS_CMD_STOP_LOGGING_REQ = 0x78 | Command to close the file
| PC to Device | [8 empty] | Response Command | M2M2_FILE_SYS_CMD_STOP_LOGGING_RESP = 0x79
[10 empty] | Status | M2M2_FILE_SYS_STATUS_LOGGING_STOPPED = 0x4D (or) | File closed successfully performed successfully
| Header | [9 empty] | M2M2_FILE_SYS_STATUS_ERROR = 0x42 (or) | Error in Closing the file
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_FILE_SYS_STATUS_LOGGING_NOT_STARTED = 0x50 | Logging not yet started
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C5 | 01 | 00 | 0B | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11
| Description | 1: Command; 2: Status; 3: stop type
| Data | 1: 78; 2: 00; 3: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C5 | 01 | C7 | 05 | 00 | 0B | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11
| Description | 1: Command; 2: Status; 3: stop type
| Data | 1: 79; 2: 41; 3: 00

## SQIStart
SQI Start Commands
[5 empty] | 1 | Load Adpd cfg |  | 4 | Sensor SQI Start |  | 8 | sub radpd6 add
[5 empty] | 2 | Set Adpd freq to 100Hz |  | 5 | Sub rsqi add
[5 empty] | 3 | ClockCalibration |  | 6 | AGC ON
[5 empty] | 4 | Set SQI Slot |  | 7 | Sensor adpd start
Load Adpd configuration | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_SENSOR_ADPD4000 = 0xC110
[13 empty] | Request Command |  | M2M2_SENSOR_ADPD_COMMAND_LOAD_CFG_REQ = 0x42 |  | Command to Load Adpd device configuration
| PC to Device | [11 empty] | Response Command |  | M2M2_SENSOR_ADPD_COMMAND_LOAD_CFG_RESP = 0x43
[13 empty] | Device ID |  | M2M2_SENSOR_ADPD4000_DEVICE_4000_G = 0x28 |  | For Green LED:M2M2_SENSOR_ADPD4000_DEVICE_4000_G = 0x28 For Red LED:M2M2_SENSOR_ADPD4000_DEVICE_4000_R = 0x29 For IR LED:M2M2_SENSOR_ADPD4000_DEVICE_4000_IR = 0x2A For Blue LED:M2M2_SENSOR_ADPD4000_DEVICE_4000_B = 0x2B For G, R, IR, B LEDs: M2M2_SENSOR_ADPD4000_DEVICE_4000_G_R_IR_B = 0x2C
| Header | [11 empty] | Status |  | M2M2_APP_COMMON_STATUS_OK = 0 (or) |  | ADPD device loaded successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | [5 empty] | M2M2_APP_COMMON_STATUS_ERROR = 1 (or) |  | Error in Loading ADPD device
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 00 | 15 | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
| Description | Command | Status | Device ID | return data | return data | return data | return data | return data | return data | return data | return data | return data | return data
| Data | 42 | 00 | 28 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 10 | C7 | 05 | 00 | 15 | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
| Description | Command | Status | Device ID | return data | return data | return data | return data | return data | return data | return data | return data | return data | return data
| Data | 43 | 00 | 28 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00
Register Write Command to set ADPD at 100Hz
PC to Device
|  | Write ADPD register 0x0D | [7 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
Header |  | 17: Destination; 19: M2M2_ADDR_SENSOR_ADPD = 0xC110
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | Request Command |  | M2M2_SENSOR_COMMON_CMD_WRITE_REG_16_REQ = 0x23 |  |  |  | Command to Perform a register operation on a device.
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB |  | Response Command |  | M2M2_SENSOR_COMMON_CMD_WRITE_REG_16_RESP = 0x24
Data | C7 | 05 | C1 | 10 | 00 | 0F | 00 | 00 |  | Status |  | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) |  |  |  | Register read/write operation performed successfully
Payload |  | 19: M2M2_APP_COMMON_STATUS_ERROR = 0x1 (or) |  | Error in Read/Write register opertaion
Byte Number | 9 | 9: 10; 10: 11; 11: 12; 12: 13; 13: 14; 14: 14
Description | Command | 9: Status; 10: Num_ops; 11: Address LSB; 12: Address MSB; 13: Value LSB; 14: Value MSB
Data | 23 | 9: 00; 10: 01; 11: 0D; 12: 00; 13: 10; 14: 27
Device to PC
Header
Byte Number | 1 | 9: 2; 10: 3; 11: 4; 12: 5; 13: 6; 14: 7; 15: 8
Description | Source MSB | 9: Source LSB; 10: Destination MSB; 11: Destination LSB; 12: Length MSB; 13: Length LSB; 14: Checksum MSB; 15: Checksum LSB
Data | C1 | 9: 10; 10: C7; 11: 05; 12: 00; 13: 0F; 14: 00; 15: 00
Payload
Byte Number | 9 | 9: 10; 10: 11; 11: 12; 12: 13; 13: 14; 14: 14
Description | Command | 9: Status; 10: Num_ops; 11: Address LSB; 12: Address MSB; 13: Value LSB; 14: Value MSB
Data | 24 | 9: 00; 10: 01; 11: 0D; 12: 00; 13: 10; 14: 27
Clock Calibration |  | 20: source | M2M2_ADDR_APP_CLI = 0xC705
|  | 20: Destination | M2M2_ADDR_SENSOR_ADPD4000 = 0xC110
|  | 20: Request Command | M2M2_SENSOR_ADPD_COMMAND_CLOCK_CAL_REQ = 0x44 |  | Calibrate the 32M and 32K clock to reduce the deviation to a minimum
| PC to Device | 20: Response Command | M2M2_SENSOR_ADPD_COMMAND_CLOCK_CAL_RESP = 0x45
|  | 20: Status | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) |  | Clock calibration done successfully
| Header | [13 empty] | M2M2_APP_COMMON_STATUS_ERROR = 0x1 (or) |  | Error in clock claibration
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  |  |  | clock cal id |  | 6-> in DVT1 Watches; 2 -> DVT2 Watches
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 00 | 0B | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11
| Description | 1: Command; 2: Status; 3: clockcalid
| Data | 1: 44; 2: 00; 3: 02
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 10 | C7 | 05 | 00 | B | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11
| Description | 1: Command; 2: Status; 3: clockcalid
| Data | 1: 44; 2: 00; 3: 02
Set SQI Slot for Slot F Green LED | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_MED_SQI = 0xC80C
[13 empty] | Request Command |  | M2M2_SQI_APP_CMD_SET_SLOT_REQ = 0x5E |  | Command to set Adpd device slot
[13 empty] | SQI Slot |  | 0x20 |  | For Slot F -> 2^5
[13 empty] | Response Command |  | M2M2_SQI_APP_CMD_SET_SLOT_RESP = 0x5F
[13 empty] | Status |  | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) |  | ADPD slot set successfully
[15 empty] | M2M2_APP_COMMON_STATUS_ERROR = 0x1 (or) |  | Error in setting ADPD slot
PC to Device
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C8 | 0C | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: SQI SLot; 4: SQI Slot
| Data | 1: 5E; 2: 00; 3: 20; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 01 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: ADPD SlotA mode; 4: ADPD SlotB mode
| Data | 1: 47; 2: 00; 3: 00; 4: 14
Start SQI sensor | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_SENSOR_ADPD = 0xC80C
[13 empty] | Request Command |  | M2M2_APP_COMMON_CMD_STREAM_START_REQ = 0x4 |  | Command to start sensor
| PC to Device | [11 empty] | Response Command |  | M2M2_APP_COMMON_CMD_STREAM_START_RESP = 0x5
[13 empty] | Status |  | M2M2_APP_COMMON_STATUS_STREAM_STARTED = 0x2 |  | Stream started successfully
| Header | [13 empty] | M2M2_APP_COMMON_STATUS_STREAM_IN_PROGRESS = 0x4 |  | Already Stream started
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | [5 empty] | M2M2_APP_COMMON_STATUS_STREAM_NOT_STARTED = 0x7 |  | Error in starting stream
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C8 | 0C | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 04; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 01 | C8 | 0C | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 05; 2: 02; 3: 00; 4: 00
Subscribe SQI Stream | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_SENSOR_MED_SQI = 0xC80C
[13 empty] | Request Command |  | M2M2_APP_COMMON_CMD_STREAM_SUBSCRIBE_REQ = 0xC |  | Command to subscribe the stream
| PC to Device | [11 empty] | Response Command |  | M2M2_APP_COMMON_CMD_STREAM_SUBSCRIBE_RESP = 0xD
[13 empty] | Stream address |  | M2M2_ADDR_SENSOR_MED_SQI_STREAM = 0xC201
| Header | [11 empty] | Status |  | M2M2_APP_COMMON_STATUS_SUBSCRIBER_ADDED = 0x9 |  | Subscriber added successfully to the stream mentioned
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C8 | 0C | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0C; 2: 00; 3: 0D; 4: C8
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 01 | C8 | 0C | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0D; 2: 09; 3: 0D; 4: C8
TurnAGC ON | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_SENSOR_ADPD = 0xC110
[13 empty] | Request Command |  | M2M2_SENSOR_ADPD_COMMAND_AGC_ON_OFF_REQ = 0x76 |  | Command to set Adpd device slot
[13 empty] | agc_type |  | 0x1 |  | To Select between green/blue/red/ir/mwl(all 4 LEDs) mwl = 0 g = 1 r = 2 ir = 3 b = 4
[13 empty] | agc_cntrl |  | 0x1 |  | To Enable AGC, agc_cntrl= 1 To Disable AGC, agc_cntrl= = 0
[13 empty] | Response Command |  | M2M2_SENSOR_ADPD_COMMAND_AGC_ON_OFF_REQ = 0x77
[13 empty] | Status |  | M2M2_APP_COMMON_STATUS_OK = 0x0 (or) |  | AGC Control set successfully
[15 empty] | M2M2_APP_COMMON_STATUS_ERROR = 0x1 (or) |  | Error in setting AGC Control
PC to Device
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 00 | 0D | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 12
| Description | Command | Status | num_of_ops | agc_type | agc_cntrl
| Data | 76 | 00 | 01 | 01 | 01
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 01 | C7 | 05 | 00 | 0D | 00 | 00
Payload
| Byte Number | 9 | 10 | 11 | 12 | 12
| Description | Command | Status | num_of_ops | agc_type | agc_cntrl
| Data | 77 | 00 | 01 | 01 | 01
Start ADPD sensor | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_SENSOR_ADPD = 0xC110
[13 empty] | Request Command |  | M2M2_APP_COMMON_CMD_STREAM_START_REQ = 0x4 |  | Command to start sensor
| PC to Device | [11 empty] | Response Command |  | M2M2_APP_COMMON_CMD_STREAM_START_RESP = 0x5
[13 empty] | Status |  | M2M2_APP_COMMON_STATUS_STREAM_STARTED = 0x2 |  | Stream started successfully
| Header | [13 empty] | M2M2_APP_COMMON_STATUS_STREAM_IN_PROGRESS = 0x4 |  | Already Stream started
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | [5 empty] | M2M2_APP_COMMON_STATUS_STREAM_NOT_STARTED = 0x7 |  | Error in starting stream
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 04; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 10 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 05; 2: 02; 3: 00; 4: 00
Subscribe ADPD4000 Green Stream | [12 empty] | source |  | M2M2_ADDR_APP_CLI = 0xC705
[13 empty] | Destination |  | M2M2_ADDR_SENSOR_ADPD_4000_G = 0xC110
[13 empty] | Request Command |  | M2M2_APP_COMMON_CMD_STREAM_SUBSCRIBE_REQ = 0xC |  | Command to subscribe the stream
| PC to Device | [11 empty] | Response Command |  | M2M2_APP_COMMON_CMD_STREAM_SUBSCRIBE_RESP = 0xD
[13 empty] | Stream address |  | M2M2_ADDR_SENSOR_ADPD_STREAM6 = 0xC216
| Header | [11 empty] | Status |  | M2M2_APP_COMMON_STATUS_SUBSCRIBER_ADDED = 0x9 |  | Subscriber added successfully to the stream mentioned
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0C; 2: 00; 3: 16; 4: C2
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 10 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0D; 2: 09; 3: 16; 4: C2

## SQIStop
SQI Stop Commands
UnSubscribe ADPD Stream | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_SENSOR_ADPD4000 = 0xC110
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_UNSUBSCRIBE_REQ = 0xE | Command to unsubscribe the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_UNSUBSCRIBE_RESP = 0xF
[10 empty] | Stream address | M2M2_ADDR_SENSOR_ADPD_STREAM6 = 0xC216
| Header | [8 empty] | Status | M2M2_APP_COMMON_STATUS_SUBSCRIBER_REMOVED = 0xA | Stream unsubscribed successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_SUBSCRIBER_COUNT_DECREMENT = 0xB | If stream subscribed more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0E; 2: 00; 3: 16; 4: C2
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 10 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0F; 2: 0A; 3: 16; 4: C2
Stop ADPD Sensor | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_SENSOR_ADPD4000 = 0xC110
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_STOP_REQ = 0x6 | Command to stop the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_STOP_RESP = 0x7
[10 empty] | Status | M2M2_APP_COMMON_STATUS_ERROR = 0x1
| Header | [9 empty] | M2M2_APP_COMMON_STATUS_STREAM_STOPPED = 0x3 | Stream stopped successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_STREAM_COUNT_DECREMENT =0x 6 | If stream started more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C1 | 10 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 06; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C1 | 10 | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 07; 2: 03; 3: 00; 4: 00
UnSubscribe SQI Stream | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_SENSOR_MED_SQI = 0xC80c
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_UNSUBSCRIBE_REQ = 0xE | Command to unsubscribe the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_UNSUBSCRIBE_RESP = 0xF
[10 empty] | Stream address | M2M2_ADDR_SENSOR_MED_SQI_STREAM = 0xC80D
| Header | [8 empty] | Status | M2M2_APP_COMMON_STATUS_SUBSCRIBER_REMOVED = 0xA | Stream unsubscribed successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_SUBSCRIBER_COUNT_DECREMENT = 0xB | If stream subscribed more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C8 | 0C | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0E; 2: 00; 3: 0D; 4: C8
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C8 | 0C | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 0F; 2: 0A; 3: 0D; 4: C8
Stop SQI Sensor | [9 empty] | source | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Destination | M2M2_ADDR_SENSOR_MED_SQI = 0xC80C
[10 empty] | Request Command | M2M2_APP_COMMON_CMD_STREAM_STOP_REQ = 0x6 | Command to stop the stream
| PC to Device | [8 empty] | Response Command | M2M2_APP_COMMON_CMD_STREAM_STOP_RESP = 0x7
[10 empty] | Status | M2M2_APP_COMMON_STATUS_ERROR = 0x1
| Header | [9 empty] | M2M2_APP_COMMON_STATUS_STREAM_STOPPED = 0x3 | Stream stopped successfully
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | M2M2_APP_COMMON_STATUS_STREAM_COUNT_DECREMENT =0x 6 | If stream started more than one time
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C7 | 05 | C8 | 0C | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 06; 2: 00; 3: 00; 4: 00
Device to PC
Header
| Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
| Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
| Data | C8 | 0C | C7 | 05 | 00 | 0C | 00 | 00
Payload
| Byte Number | 1: 9; 2: 10; 3: 11; 4: 12
| Description | 1: Command; 2: Status; 3: Stream Address LSB; 4: Stream Address MSB
| Data | 1: 07; 2: 03; 3: 00; 4: 00

## SQIData
Sample SQI data from CLI
c8:0d:c7:05:00:19:00:00:28:00:72:02:00:00:00:01:00:20:00:00:95:23:F8:19:00
SQI Data Stream | [9 empty] | source | M2M2_ADDR_MED_SQI_STREAM = 0xC80D
[10 empty] | Destination | M2M2_ADDR_APP_CLI = 0xC705
[10 empty] | Command | M2M2_SENSOR_COMMON_CMD_STREAM_DATA = 0x28
Device to PC
Header
Byte Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Description | Source MSB | Source LSB | Destination MSB | Destination LSB | Length MSB | Length LSB | Checksum MSB | Checksum LSB
Data | C8 | 0D | C7 | 05 | 00 | 19 | 00 | 00
Payload
Byte Number | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18
Description | Command | Status | Sequence Number LSB | Sequence Number MSB | sqi LSB | sqi | sqi | sqi MSB | nSQISlot LSB | nSQISlot MSB
Data | 28 | 00 | 72 | 02 | 00 | 00 | 00 | 01 | 00 | 20
Payload
Byte Number | 19 | 20 | 21 | 22 | 23 | 24 | 25
Description | nAlgoStatus LSB | nAlgoStatus MSB | TimeStamp LSB | TimeStamp | TimeStamp | TimeStamp MSB | Reserved
Data | 00 | 00 | 95 | 23 | F8 | 19 | 00